## Sheet1
Construction of Hadron Treatment and Support Facility at ACTREC Campus, Tata Memorial Centre on Design Build Basis
Minutes of PREBID Meeting held at 10.30 AM on 29.04.2016 at TATA MEMORIAL HOSPITAL, PAREL, MUMBAI - Reply to second set of queries
SL. NO. | REFERENCE SECTION /CLAUSE OF BID DOCUMENT | PROVISION IN BID DOCUMENT | PRE-BID QUERY FROM BIDDERS | CLARIFICATION
1 | RFP Vol.2, Pg.54 - Clause 18.4 | Consequences of Employer's Risks If as a result of rectifying this loss or damage, there is any delay in the performance of the Contract, the EPC Developer shall give a further notice to the Engineer and shall be entitled subject to an extension of time for any such delay, if completion is or will be delayed under Sub-Clause 9.4 (Extension of Time for Completion). No extra payment shall be given to the EPC Developer for any delay suffered due to the above reason. | We request you to modify the clause to include cost compensation towards delays caused by Employer's Risks. | As per tender document.
2 | RFP Vol.4 - Clause 9.3 | The source of water supply for the proposed buildings shall be Bore wells. The Bore wells location shall be considered inside the campus with the permission of competent authority. However the number of bore wells shall be based on the yielding capacity of bore well. | While site visit it was informed that, there is existing municipal water supply connection, we need to just transfer the water from existing underground tank to proposed dedicated u/g tank of proposed building. But, as per RFP, source of water supply shall be Bore wells. Kindly clarify the contradiction about the source of water. | Refer point no.118 of our pre-bid reply dated 25.05.2016
3 |  | Material of construction | We are proposing below mentioned MOC, Please confirm a.Hydrant & Sprinkler Pipe - M.S. Class 'C' b.Landing Valves - G.M. | Agreed subject to their confirming to latest IS-codes.
4 |  | Supression system | Presently we have not considered any clean agent supression system / gas based supression system in our scope of work. Please confirm. | Refer point no.183 of our pre-bid reply dated 25.05.2016
5 |  | Scope Clarification | In received site plan, the compressor room is shown near to the proposed building. We are understanding that, there is existing medical gas bank/ manifold , compressor, vacuum pump near to the proposed building, we just need to take connection from those existing medical gas plant (i.e. high side system is not in our scope of work). Please confirm. | Refer point no.141 of our pre-bid reply dated 25.05.2016
6 |  | Material of construction | Pipe Materials to be used conforming to HTM 2022 standards / Lloyd certification from reputed manufacturer of medical gas copper pipes. Please confirm. | Copper Pipe with Lloyd certification comprising of diameter and thickness as per HTM 2022 to suit the design.Isolated controls, alarms etc to be provided as per HTM 2022
7 |  | Gas Requirements | Please share the area wise medical gas requirement (i.e. type of gas and no. of outlets) | Refer reply to query No.23
8 |  | AGSS | Is there any anesthesia gas scavanging system requirement (AGSS). Kindly specify. | AGSS system should comply with HTM 2022.It should be with suction motor installed at a distance but invariably controllable from each treatment room. Suitable isolation arrangement should be provided to eliminate the backflow of anaesthesia gas into any of the treatment room.
9 |  | Make List | Please share the approved make list of medical gas system, If available. | All components shall confirm to HTM 2022
10 |  | CHILLED WATER TAPPING TO THE WATER COOLING SYSTEM | We understand that the chilled water tapping to the water cooling system to be taken from the existing hospital chillers, is in SPCL scope.Kindly confirm the same and also share the details like chilled water quantity or pipe sizes,existing hospital chiller location,etc so as to account the chilled water tapping piping length. | Connection to water cooling system is not from the Existing Hospital Chillers
11 |  |  | Any kind of fixed, loose furniture is not in our scope. | Confirmed
12 |  | GLASS PROPERTIES | Kindly provide the thermal property details ( i.e U value and SHGF) of glass to be considered for glazing purpose. | EPC developer to propose suitable glass which will be confirmed during design review.
13 |  |  | Is there any specific type of water requirement for the proposed building ? Otherwise we are considering only potable domestic water everywhere (even inside the , cyclotron manifolds, any specific area etc.) and water shall be supplied by gravity system. | Please refer IV.Cooling and Gasses 51.01B of revised IBD document.
14 |  |  | Are we going to use the STP (recycle water for coolong tower (HVAC make up) please specifiy the required hardness in ppm | It is proposed to install Air cooled Chillers at Hadron. Hence no make up water for Cooling Towers
15 | IBD_Latest DOC_1. Cyclo Vault Layout |  | In the received cyclo vault layout, there are three water manifold and one compressed air manifold shown inside the cyclotron vault. We are understanding that we need to provide the water and compressed air upto the manifold, i.e. the manifolds and the distribution network after manifold is not in our scope of work. Please confirm. | Shall be as per the revised IBD & to be confirmed by IBA
16 | IBD_Latest DOC_2. Gantry Treatment room | Floor shall have a slight slope to accommodate drainage to sumps | We are understanding the drainage sump with sump pump in each gantry pit (total three numbers) to be provided. Kindly confirm. | Yes
17 | IBD_Latest DOC_2. Water cooling room |  | Kindly share the total water requirement (per day & per hour) for makeup of cooling plant in water cooling room and also specify the type of water requirement. Otherwise we are considering potable domestic water which shall be supplied by gravity system. | Please refer IV.Cooling and Gasses 51.01B, 51.02B& 51.03B of revised IBD document.
18 | IBD_Latest DOC_3. Machine Shop | 3 Compressed air outlets at each milling machine + 1 additional. 4.Industrial size double sink with cold and warm water. 5.Cold water outlet at each milling machine. 7. Floor drain at each milling machine | As we have to provide utility for machine room please specify the locations of machine rooms | Machine room for utilities may be located appropriately by the EPC Developer in their design proposal.
19 | IBD_Latest DOC_3. Maintenance Room | 1.3 A sink with cold and hot water 1.5. Compressed air outlets | We are unable to find any room named as "Maintenance room". Kindly share the layout for the same | Please refer the revised drawings shared on 14th May 2016
20 | IBD_Latest DOC_3. | IV_Cooling and Gases_Cooling water schematic distribution | We are presently considering only the make up water supply upto the cooling room, i.e. (make up water any distribution network from the cooling room (supply & return pipe) to the equipements and different strategic locations are not considered in our scope of work. Please confirm. | Please refer IV.Cooling and Gasses 51.02B of revised IBD document.
21 | IBD_Latest DOC_3. IV_Cooling and Gases_Power supply room water distribution | 4 Pipes (Supply & Return) shall pass through the walls separating the water-cooling room and the power supply room. If the water pipes are embedded in the concrete, then 4 additional spare pipes shall be foreseen. | Presently we are not considering any water supply and return (4 pipes) inside power supply room in our scope of work. Please confirm. | As per revised IBD document, 51.02B & 51.13B, it is under the scope of the EPC Developer
FIRE FIGHTING
22 | IBD_Latest DOC_2. Gantry Treatment room_Treatment area | All rooms where patient are present will be equipped with fire detectors and water sprinklers in compliance withh applicable regulationss the treatment rooms. This requirement shall be specified by the client/user | As Per NBC 2005, Table 23, automatic sprinklers are not required for proposed building. Hence, we are presently not considering any sprinklers for building. Kindly confirm. | Sprinkler system is to be proposed in the supporting facility and for Hadron Facility, inergen gas system is to be proposed.
MEDICAL GAS
23 | IBD_Latest DOC_2. Gantry Treatment room_Treatment area | PTEV recommends D/BT to supply and install a sufficienamount of medical gas outlets in the treatment rooms. This requirement shall be specified by the client/user | Kindly specify the type of gases and no. of outlets required in treatment rooms (TCRs) | Each treatment room should have 2 oxygen outlets,2 medical gas outlets, 2 vaccuum outlets and 1 medical air outlet on the wall through standard female sockets (HTM 2022). Nitrous oxide outlet if available shall be provided one in each room.In absence of nitrous oxide, anaesthesia machines shall be built in replaceable N20 Cylinder.
24 |  |  | We are understanding the high side part of medical gas system (such as manifold, compressor, vacuum plant etc) is not in our scope of work. Only we are considering the the pipe routing from the manifold to the strategic locations. | Refer point no.139, 140, 141 of our pre-bid reply dated 25.05.2016
25 |  |  | Please provide the tentative locations of gas supply and also the type of gases, no. of outlets.
26 |  |  | We are considering the routing upto to desired rooms only. The detailed routing from that location to the equipments are not considered in our scope of work. Please confirm.
27 |  |  | Please give all IBD drawing in Auto cadd format and finishes schedule for proton therapy facility. | IBD drawing in Auto cadd format cannot be shared. The finishes schedule for the proton Therapy facility to be proposed by the EPC Developer for scrutiny
28 | IBD Doc 1- II-Rooms Cyclotron/ ESS/ BTS vaults |  | Please give details/specifictaions of Moisture detection system, drainage system, Fencing, removable shielding blocks masonary, monorail with chain hoist, water and electrical cable trench. | Humidity sensors shall be installed to detect the moisture level. Moisture to be controlled by duct heaters installed with air conditioning system. Level of humidity (40%-50%) subject to the requirement. Refer 31.01 B of revised IBD
29 | IBD Doc 1- II-Rooms Cyclotron/ESS/BTS vaults- 1.11- 2 A 2/2 |  | Tables given is not readable. Please give readable copy | Please refer the revised IBD shared on 14.05.2016
30 | Doc 2- II-Rooms-Gantry treatment rooms- 2.11, 2.13 A |  | Table given is not readable. Please give readable copy | Please refer the revised IBD shared on 14.05.2016
31 | Doc 2- II-Rooms-Gantry treatment rooms- 2.31A |  | Please give Internal wall finishes detail/specifictaions | Refer: 02.32.B - Revised IBD-Part1
32 | Doc 2- II-Rooms-Gantry treatment rooms- 2.32A |  | Please give details/specifictaions for door, ladder, adjustable convex mirror, medical gas outlet, walk on grating, metallic staircase. | EPC Bidder to propose & will be confirmed during design review.
33 | Doc 2- II-Rooms-Gantry treatment rooms- 2.33-1A |  | Please give details/specifictaions for Jib crane in Gantry area. | Refer : 02.33-1 .B - Revised IBD-Part1
34 | Doc 2- II-Rooms-Gantry treatment rooms- 2.51 A |  | Please give details/specifictaions for Moisture detection system | Refer reply to query No.28 above
35 | Doc 2- II-Rooms-Gantry treatment rooms- 2.52 A |  | Please give details/specificaions for Embedded anchor plates for monorail, false ceiling | Refer page no: 01.32.2.B IBD-REV-B
36 | Doc 2- II-Rooms-Gantry treatment rooms- 2.56 A |  | Please give details/specifictaions for protection painting coat | Refer page no: 00.00B IBD-REV-B
37 | Doc 2- II-Rooms-Gantry treatment rooms- 2.75 A |  | Please give details/specifictaions for Sink with hot & cold water | EPC Bidder to propose & will be confirmed during design review.
38 | Doc 2- II-Rooms-Power supply rooms- 31.01A |  | Please give details/specifictaions for Overhead monorail beam, wall finish concrete or metallic having shield of electromagnetic radiation | EPC Bidder to propose & will be confirmed during design review.
39 | Doc 2- II-Rooms-Power supply rooms- 31.01A |  | As per details given room size may be modified depending on requirement. Please specify exact room size with detailed plan | Please refer the revised drawings shared on 14.05.2016
40 | Doc 2- II-Rooms-Power supply rooms- 31.01A |  | Please give exact scope of D&B team | As per IBD-REV-B shared on 14.05.16
41 | Doc 2- II-Rooms-Power supply rooms- 31.01A |  | Please give details/specifictaions for Sink with hot & cold water, Antistatic nonconductive raised floor. | EPC Bidder to propose & will be confirmed during design review.
42 | Doc 2- II-Rooms-Water cooling rooms- 32.01A |  | Tables given are not readable. Please give readable copy | Please refer the revised IBA REV-B documents shared on 14.05.2016
43 | Doc 2- II-Rooms-Water cooling rooms- 32.01A |  | Please give details/specifictaions for Door details, Concrete pad, wall between cooling room & power supply room, finish details to achieve noise control in walls | EPC Bidder to propose & will be confirmed during design review.
44 | Doc 2- II-Rooms-Main control room-33.01A |  | Table 1 is not readable. Please give readable copy | Please refer the revised IBA REV-B documents shared on 14.05.2016
45 | Doc 2- II-Rooms-Main control room-33.01A |  | Please give details/specifictaions for Antistatic raised floor, moisture detection system, suspended dismountable ceiling, cupboard | EPC Bidder to propose & will be confirmed during design review.
46 | Doc 2- II-Rooms-Treatment control rooms- |  | Room image is not visible. | Please refer the revised IBA REV-B documents shared on 14.05.2016
47 | Doc 2- II-Rooms-Treatment control rooms- 34.01A AND Doc 2- II-Rooms-PTEV server Room |  | Please give details/specifications for customer equipment & furniture, cable chase for computers, ceiling with indirect lighting with dimmer control above wall | The designer may propose the details and specifications approximately.
48 | Doc 2- II-Rooms-Machine shop | This room is optional | Please specify if this room is required or not. If yes then give details/plans. | Machine room for utilities may be located appropriately by the EPC Developer in their design proposal.
49 | Doc 3- II-Rooms-Other rooms |  | Please specify location and scope of maintenance room, PTEV spare parts storage room, PTEV offices | Please refer the revised drawings shared on 14.05.2016
50 | Doc 3- II-Rooms-Recommended rooms |  | Please specify if these room is required or not. If yes then give details/plans. Treatment planning computer room (tpcr), Tps/ois server room, patient immobilization devices storage room, patient treatment devices storage room, Patient treatment devices controlled storage room, patient immobilization body devices fabrication shop, Dosimetry equipment storage room, snout storage room, processing laboratory | Please refer the revised drawings shared on 14.05.2016
51 | Doc 3- III-INSTALLATION-50.01-1 A |  | Table 1, 2 is not readable. Please give readable copy | Please refer the revised IBA REV-B documents shared on 14.05.2016
52 | Doc 3- III-INSTALLATION-50.01-1 A |  | Please give details/specifictaions for Watertight permanent hatch panel, fencing details for storage equipment | Specifications/design specifications of material may be submitted by EPC Developer for scrutiny
53 | Doc 3- III-INSTALLATION-50.01-2 A |  | Please give details/specifictaions for waterproof rigid enclosure for equipments, high strength grout | Specifications/design specifications of material may be submitted by EPC Developer for scrutiny
54 | Doc 3- IV-COOLING & GASES- 51.03- A |  | Table 1, 2 is not readable. Please give readable copy | Please refer the revised IBA REV-B documents shared on 14.05.2016
55 | Doc 3- IV-COOLING & GASES- 51.11- A |  | Please give details/specifictaions for metal gratings Table 1, 2 is not readable | Specifications/design specifications of material may be submitted by EPC Developer for scrutiny
56 | Doc 3- IV-COOLING & GASES- 51.22- A |  | Table is not readable. Please give readable copy | Please refer the revised IBA REV-B documents shared on 14.05.2016
57 | Doc 3- V-ELECTRICAL SET-52.11- 2 A |  | Table is not readable. Please give readable copy | Please refer the revised IBD shared on 14.05.2016
58 | Doc 5- V-OTHER TECHNICS- 61.01 A |  | Tables is not readable. Please give readable copy | Please refer the revised IBD shared on 14.05.2016
59 | Doc 5- V-OTHER TECHNICS- 62.01 A | Client has proposed to refer doc "Radiation Sources Proteus 235" (MID 35597)" for shielding from radiations | Please give details/specifications for shielding to be provided against radiations | Specifications/design specifications for shielding to be proposed by EPC Developer as per AERB norms and this will be confirmed by IBA/HLL during design review.
60 | Doc 5- V-OTHER TECHNICS- 62.01 A | The Customer may want to consider the need for a storage area for radioactive parts. D/BT will put a “closet” or Rated Cabinet for storage of radioactive parts against the ESS shielding wall. This closet shall be large enough to contain a four-drawer file cabinet plus any user needs. This storage room is out of PTEV scope and up to the Customer. | Please specify if this radiation closet is required or not. If yes then give details/plans. | Radiation sheilding proposal to be submitted by D/BT which will be confirmed by IBA/HLL during design review.
61 | Doc 5- V-OTHER TECHNICS- 62.01 A | 9.1. Refer to figure 38.01: Patient Treatment Devices controlled storage room . 9.2. This storage room is out of PTEV scope and up to the Customer. | Please specify if this area is required or not. If yes then give details/plans. | Please refer the revised drawings shared on 14.05.2016
62 | Doc 5- V-OTHER TECHNICS- 62.01 A | 10.2. PTEV recommends that the D/BT fill the gaps with borated foam after the installation of the concrete blocks to avoid any radiation leakage. | Please give details/specifications for borated foam to avoid radiation leakage | Refer point no.83 of our pre-bid reply dated 25.05.2016
63 | Doc 5- VI-OTHER TECHNICS- THERAPY SAFETY SYSTEM (TSS) 63.01 A | The TSS components locations and system interface and requirements are shown on this figure and on Figure 63.02-Safety System layout and Figure 63.05-TSS Table of Operands. | Fig. Figure 63.02-Safety System layout and Figure 63.05-TSS Table of Operands are not readable. Please give readable copy. Please specify exact D&B T scope. | Please refer the revised IBD shared on 14.05.2016
64 | Doc 5- VI-OTHER TECHNICS- (TSS) 64.01 A | The system for the detection and monitoring of ionizing radiation is the responsibility of the D/BT. | Please give details/specifications for detection and monitoring of ionizing radiation. | D/BT may propose an appropriate system
65 | Doc 5- VI-OTHER TECHNICS- (TSS) 68.01 A | The system for communication and surveillance is the responsibility of the D/BT and shall be in accordance with the Customer needs. | Please give details/specifications for communication equipments as PTEV has give tentative requirement for the same. Table is not readable. Please give readable copy | Please refer the revised IBD shared on 14.05.2016
66 |  |  | Please specify Battery Limit for Gantry Area (only isolator to be provided?) | Circuit breaker of suitable rating has to provided at the various hook up postions
67 |  |  | Please confirm any Special light fixture, Power socket or other electrical requirement in Gantry area (requirement of - IP/ Flame proof fixture/ socket etc) | For lighting refer IBD doc. 66.01 B / IBA
68 |  |  | Please confirm RJ-11 socket , 2 Pair 0.5 mm tele. Cu. Wire to be used i.e. Analogue type telephone distribution is considered | RJ-45 socket and 4 pair Cat 5e cable. Please refer RFP.
Reply to the query which were not answered in the minutes of pre bid meeting uploaded in website on 25.05.2016
163 |  |  | Power failure backup for DG sets shall be stand by duty | Power failure backup for DG sets shall be stand by
173 |  |  | Active components like switches, servers, routers, software etc., for data networking & EPABX for Telephone networking are excluded from the Bidders scope of works. Kindly confirm. | EPABX for Telephone networking is included in the scope of EPC Developer.
176 |  |  | All Special systems related to Proton Therapy like 1. Therapy Safety System 2. Radiation Detection System 3. Medical Vision Cameras inside the Gantry 4. Moisture Detection Systems are excluded. Confirm. | Therapy Safety System,Medical Vision Cameras inside the Gantry,Moisture Detection Systems are included in the scope of EPC Developer.
177 |  |  | All ELV/IT Systems like 1. Public Address System 2. Building Management System 3. Access Control System 4. Hospital Management System 5. Queuing Management System 6. Audio Visual System 7. Picture archieving Communication Systems are excluded. Confirm. | Public Address System, Access Control System are included in the scope of EPC Developer.
179 |  |  | We are not considering any active components viz. Core switch, layer switch, access switch, WiFi router, Optic fibre cable, etc. and any other IT networking components and Severs, PC workstations, etc. | 2 Nos of 12 core OF cable to be laid to the proposed building from: a) PS Building & b) JS building. Racks to be provided as per IBD document (67.01 B) - Network topology and RFP
181 |  |  | We are considering only Addressable Fire Alarm System, Telephone distribution system, Data distribution system & IP CCTV system and no other LV system is considered. | ELV systems in the scope of EPC Developer: 1. Telephone System 2. Data System 3. Fire Alarm System 4. PA System 5. CCTV System 6. Nurse Call System 7. CATV System 8. Therapy Safety System (For Hadron Treatment Room) 9. Moisture Detection System (For Hadron Treatment Room)
GENERAL
Further queries will not be entertained.