Corporate and Registered Office,

HLL Bhavan, Poojappura,

Thiruvananthapuram– 695 012

Kerala, India.

Phone: 0471 – 2354949

Website: www.lifecarehll.com

## Invitation for Bids

Battery operated Power drill and Power Saw

Date	:		19.08.2015

IFB No.	:		HLL/CHO/HCS/PROC/2015-16/279/TEN-13

The schedule of the bid is given below.

Last date and time for receipt of bids  :  15.00 Hrs on 07.09.2015

Time and date of opening of Bids         :  15.30 Hrs on 07.09.2015

###### Dear Sir,

Sub: Supply Battery operated Power Drill & Power Saw.

HLL Lifecare Limited, Thiruvananthapuram now invites the competitive bids from the eligible bidders for supply of Battery operated power drill & Power saw.

The terms and conditions of the bid are given below.

- PRODUCT SPECIFICATIONS:

Brand name and type of Product being offered by the Firm must be clearly mentioned in the offer. Tenderer should confirm that the stores offered conform strictly to relevant specifications. Deviations, if any, are to be clearly mentioned. Complete product specifications, technical details, illustrations, literature, printed pamphlets/leaflets etc and Manufacturers price list, if any, must accompany the quotation.  Leaflets, pamphlets, etc. shall be taken for information purpose only and shall not form the part of the contract. The specifications of the required drill are given in Annexure-1.

- Prices quoted should be ‘Firm & final’ for free delivery at the sites mentioned in this tender document, mentioning the quantity, unit price, total amount and Applicable taxes etc clearly as per the price schedule enclosed (Annexure-2)

- The prices quoted shall be valid for a period of 90 days from the date of opening of bids.

- The bids will be evaluated by taking the total amount quoted for all the items in the bid.

- The bids should contain the complete technical specifications of the equipment supported with illustrative literatures.

- The items should be delivered within 3 Weeks from the date of placement of order.

- Penalty @ 0.5% per week’s delay subject to a maximum of 5% is applicable for delayed delivery.

- The bid will be summarily rejected in case any or all of the following;

      - The bid with conditional and ambiguous clauses

      - The bid without EMD

- Successful bidder should provide replacement warranty and free service of Power drill & power saw for a minimum period of one year from the date of installation. All the defective items/parts should be replaced at free of cost during the warranty period.

- EARNEST MONEY DEPOSIT (EMD)

a) The Bid should be submitted along with an Earnest Money Deposit (EMD) of INR.10,000 ( Rupees Ten thousand  Only) in the form of a Demand Draft drawn in favour of “HLL LIFECARE LIMITED, THIRUVANANTHAPURAM” of any Nationalized/Scheduled bank payable at Thiruvananthapuram. Payment in any other form will not be accepted. Bid submitted without EMD shall be summarily rejected. No interest shall be paid on any of these deposits.

b) The EMD of the successful Bidder will be treated as Security Deposit and it will be returned after the completion of the warranty period.

c) The EMD of unsuccessful bidders will be returned as promptly as possible to them within one month of selection of successful bidder.

d) The EMD will be forfeited if;

i. The Bidder withdraws the bid during the period of bid validity.

or

ii. The Bidder fails to accept the Purchaser’s corrections of arithmetic

errors in the Bidder’s bid (if any),

or

(iii). The Successful Bidder fails to accomplish the task.

- Payment will be released within 30 days from the date of Delivery.  For claiming the payment, the following documents are to be submitted.

  - Three copies of Invoice

  - Delivery report duly signed by the authorized person of HLL and representatives of the supplier.

- Charges for CAMC/AMC would be paid on half yearly basis after completion of work and satisfactory working report.

- Payment Procedure: Where there is a statutory requirement for tax deduction at source, such deduction   towards income tax and other tax as applicable will be made from the bills payable to the Supplier at rates as notified from time to time.

- Transit Insurance: Rates quoted being door delivery basis, the Supplier shall be fully responsible till full material is received in good condition at consignee’s site. As such the Supplier shall despatch the material duly insured.

- The bidder should have service support Center at Trivandrum region. The details of the same may be furnished along with the bid.

- WARRANTY:

a. Products to be offered shall be warranted for a period of One Year On Site from the date of installation. No conditional warranty shall be acceptable.

b. The Supplier shall be responsible to replace the material free of cost at site in whole or in part if found defective in any respect after receipt at site or during normal & proper usage or storage/maintenance for which the Purchaser shall give prompt written notice. Such replacements shall be effected by the Supplier within a reasonable time actually required to do so which in no case shall be more than 10 days.

- The above provisions shall also equally apply to the material replaced by the Supplier under this Clause, in case the same is again found to be defective after its replacement.

- If the Supplier fails to act with requisite promptness and thereby entails avoidable loss to the purchaser/consignee, it shall be liable to suitable action as deemed fit during the operative Warranty period.

- Charges for AMC after warrantee for 5 years shall be quoted separately as per price schedule which is attached (Annexure – 3). Party shall also submit the rate of spares separately for 5 year consumption and the rates shall be valid for 5 years.

- Charges for CAMC after warrantee shall be quoted separately as per price schedule which is attached (Annexure - 4.)

- Either AMC or CAMC will be considered after the warranty period and it will be finalized before placing the purchase order.

- The bids should be submitted at the following address.

Deputy General Manager (SD)

HLL Lifecare Limited,

Corporate and Registered Office,

HLL Bhavan, Poojappura,

Thiruvananthapuram – 695 012

- Bids should be submitted latest by 15:00 Hrs on 07.09.2015 and the same will be opened at 15:30 Hrs on the same day at the above mentioned address in the presence of the representative of the bidder who choose to attend.

- If the bid opening day is declared as holiday for HLL, the bid will be opened at the next working day of HLL.

- Any bid received after the deadline will be rejected.

- HLL reserves the right to accept or reject any or all of the bids without assigning any reason whatsoever.

- Any dispute arising out of the tender/bid document/ evaluation of bids/issue of purchase order shall be subject to the jurisdiction of the competent court at Thiruvananthapuram only.

- The envelopes containing the bid shall be bearing the words “DO NOT OPEN BEFORE …….. “ (Here insert the time and date of bid opening).

- No email or fax bids will be accepted

Thanking you,

Yours faithfully,

Smitha.L.G

Deputy General Manager (SD)

Annexure 1.

Technical Specification for Battery operated Power drill and Power Saw

| Name of Equipment | Specification |
| --- | --- |
| BATTERY OPERATED POWER DRILL AND POWER SAW | Drill & Reamer Hand Piece Selection of Drilling and reaming with built in switch option DRILL/REAM in same hand piece Selection of Drilling and reaming with same attachment. Should have dual trigger for forward/reverse and oscillation mode Maximum speed of 1200 rpm in drilling, 270 rpm in reaming Should have variable speed control on the hand piece Should deliver maximum torque of 150 in/lbs Drill torque should be 35 in/lbs Should have DC brush less motor for low maintenance With appropriate adaptors for drilling, reaming and pin placement and wire placement Future up gradation compatible for Navigation interface for Joint replacement surgeries Micro processor controlled Hand piece Can be calibrate for the consistence performance Weight of hand piece with battery should be not more than 3.5kg Fully Cannulated 4.0 mm hand piece Should have Pistol grip Hand piece Should be autoclavable Dedicated Forward and Reveres switch with safe mode Can be calibrating for the consistence performance Sagital Saw Hand piece: Should have two speed controls with standard and fast mode. Free speed of 10000- 12000 cycles per minute Micro processor controlled Hand piece Can be calibrate for the consistence performance Saw Noise level should not more than 89db Weight of hand piece with battery should be not more than 3.5 lbs Blade mount should be adjustable to different angles with 360 degree rotation Should have DC brush less motor Should be autoclavable Should have safe mode Reciprocating Saw Hand piece: Should have Safe Mode and should have minimum 13500 CPM Weight of hand piece with battery should be not more than 3.5 lbs Micro processor controlled Hand piece Can be calibrate for the consistence performance Should have DC brush less motor for low maintenance. Should have Pistol grip Hand piece Should have tool less mounting of accessories for all blades or attachments Saw noise level should not more than 93db Should be autoclavable With different blades it should have maximum speed of 13500CPM Drill and reaming Attachments: 1/4 inch Jacobs Drill Attachment with key Keyless Chuck Reamer Attachment, Pin Collet Attachment Hudson Modified Trinkle attachment K Wire Collet Attachment Battery Charger 220-240 volts charger and should have the feature to count the charging cycle for a particular battery Should have capability to identify the worn out battery Should have to charge four batteries at a time Should have an indicator to provide battery status for charging. Should be able to check over autoclaved battery cycles (Number of Time and Total time) Battery Kit: NiMh/Ni Cd batteries – 4 Should have a run time of minimum 17 minutes Should include Autoclavable outer housing Shield to protect battery from the housing Opening of battery housing for easy insertion of battery Should have option for autoclavable batteries Sterilization Case: Should be accommodate all hand piece, attachment and accessories for autoclave |

ANNEXURE-2

PRICE SCHEDULED FOR DOMESTIC GOODS OR GOODS OF FOREIGN ORIGIN LOCATED WITHIN INDIA.

| 1 | 2 | 3 | 4 |
| --- | --- | --- | --- |
|  |  |  | Price per unit (Rs) |
| Sl No | Brief description of goods | Qty nos | Rate (A) | Excise duty(if any)% and Value (B) | Sales tax/VAT/Entry tax (if any)and % value (C) | Unit Price (at consignee site basis) |

Total quoted price in Rs:-…………………………………………

In Words:-…………………………………………………………..............................................................……..

Place:-

Date:-

Name:-…………………………………………..

Business Address:-…………………………………………..

Signature of Bidder:-…………………………………………..

Seal of the Bidder:-…………………………………………..

ANNEXURE-3

ANNUAL MAINTENANCE CONTRACT PRICES SCHEDULE

| Sl No | Item description | 1st year | 2nd year | 3rd year | 4th year | 5th year | Annual maintenance contract over a period of 5 years after expiry of warranty period of 5 years from date of successful installation. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | B | C | D | E | F | G | H=C+D+E+F+G |
| 1 | Name of Equipment: Make: Model: Qty: |  |  |  |  |  |  |
| 2 | Name of Equipment: Make: Model: Qty: |  |  |  |  |  |  |

Seal and Signature of the bidder

ANNEXURE-4

COMPREHENSIVE ANNUAL MAINTENANCE CONTRACT PRICES SCHEDULE

| Sl No. | Item description | 1st year | 2nd year | 3rd year | 4th year | 5th year | Total comprehensive annual maintenance contract over a period of 5 years after expiry of warranty period of 5 years from date of successful installation. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | B | C | D | E | F | G | H=C+D+E+F+G |
| 1 | Name of Equipment: Make: Model: Qty: |  |  |  |  |  |  |
| 2 | Name of Equipment: Make: Model: Qty: |  |  |  |  |  |  |

Seal and Signature of the bidder