## TENDER FOR SUPPLY OF HOSPITAL EQUIPMENTS ## FOR ## HIMACHAL PRADESH PROJECT-II ### IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV ##### HLL LIFECARE LIMITED (AGOVT. OFINDIAENTERPRISE) AKKULAM FACTORY, SREEKARIAM P.O. THIRUVANANTHAPURAM – 17 Phone +91 471 244 5930, Fax +91 471 244 5935 Email: tradingpaft@lifecarehll.com Web: www.lifecarehll.com

##### HLL LIFECARE LIMITED

(Formerly Hindustan Latex Limited)

AKKULAM  FACTORY,   SREEKARIAM  P.O.  THIRUVANANTHAPURAM - 17

## Tender for Supply of HOSPITAL EQUIPMENTS

BID REFERENCE	       : IFB NO: HLL/AFT/TDG/HP-II/EQUIP/2014-15/IV						                     Dt.07/05/2014

DATE OF COMMENCEMENT

OF SALE OF BIDDING DOCUMENT	:	07/05/2014

LAST DATE FOR SALE OF BIDDING

DOCUMENT					:	26/05/2014   ,     	11.30 HRS

LAST DATE AND TIME FOR

RECEIPT OF BIDS				:	26/05/2014   ,     	14.00 HRS

TIME AND DATE OF OPENING

OF BIDS					:	26/05/2014   ,     	14.30 HRS

PLACE OF OPENING OF BIDS		:	HLL LIFECARE LIMITED

AKKULAM  FACTORY,

SREEKARIAM  P.O.

THIRUVANANTHAPURAM – 17

Phone +91 471 244 5930

Fax      +91 471 244 5935

Email: tradingpaft@lifecarehll.com

ADDRESS FOR COMMUNICATION		:	Joint General Manager (Mtls)

HLL LIFECARE LIMITED

AKKULAM  FACTORY,

SREEKARIAM  P.O.

THIRUVANANTHAPURAM – 17

Phone +91 471 244 5930

Fax      +91 471 244 5935

Email: tradingpaft@lifecarehll.com

ABSTRACT

SECTION No.		DESCRIPTION	                                  PAGE No.

PART  A

SECTION I		INVITATION FOR BIDS (IFB)				    04

SECTION II		INSTRUCTION TO BIDDERS (ITB)				    08

SECTION III		GENERAL CONDITIONS OF CONTRACT (GCC)		    18

SECTION IV		SPECIAL CONDITIONS OF CONTRACT (SCC)	                24

SECTION V		TECHNICAL SPECIFICATIONS 				    28

SECTION VI		QUALIFICATION CRITERIA					    62

SECTION VII		PROFORMA FOR EQUIPMENT AND QUALITY CONTROL   63

SECTION VIII		PERFORMANCE STATEMENT				    64

PART : B

SECTION IX		PRICE BID FORM						    67

SECTION X		PRICE SCHEDULE (I)					    68

SECTION XI		AMC PRICE SCHEDULE (I)					    76

SECTION XII		CHECK LIST							    89

Section XIII      Bank Guarantee Form for EMD                                                     77

Section XIV     Manufacturer‟s Authorisation Form                                                78

Section XV       Bank Guarantee Form for Performance Security /CMC Security  79

(CMC / AMC Security shall be notified by a Notary Public)

Section XVI        Proforma of Consignee Receipt Certificate  (Issue Note)           82

Section XVII      Proforma of Final Acceptance Certificate by the Consignee        83

Section XVIII     Consignee details  ----------   Annexure Attached

## PART - A

SECTION I

INVITATION FOR BIDS (IFB)

INVITATION FOR BIDS (IFB)

IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV

HLL Lifecare Limited hereby invites sealed bids from eligible  bidders for  supply, installation and commissioning of following  HOSPITAL EQUIPMENTS on tender wise, turn key contract basis for HIMACHAL PRADESH PROJECT-II, under NRHM, at District wise designated location (s), with in Himachal Pradesh.

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-01 |  |  |  |
| --- | --- | --- | --- |
| Neonate Critical Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs:) |
| 1 | Open care system: Radiant warmer, fixed height, with trolley, drawers, stand infusion double hook two poles attached. | 85 | 50000.00 |
| 2 | Phototherapy unit, Double Surface, tachable , LED Type | 12 |  |
| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-02 |  |  |  |
| Newborn Unit Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs:) |
| 1 | Resuscitator, hand-operated, neonate, 250 ml | 12 | 30000.00 |
| 2 | Resuscitator, hand-operated, neonate, 500ml | 63 |  |
| 3 | Oxygen hood S and M set of 3 each, including connecting tubes | 17 |  |
| 4 | Moniter, Vital sign, NIBP, HR, SpO2, ECG, RR and Temp. | 11 |  |
| 5 | Pulse oximeter, bedside, neonatal | 8 |  |
| 6 | Syringe pump, 10,20,50 ml, single phase | 23 |  |
| 7 | ECG unit, 3 channal, portable/SET | 6 |  |
| 8 | Oxygen concentrator | 1 |  |
| 9 | Pump, suction, Foot operated | 40 |  |
| 10 | Brest Pump | 6 |  |
| 11 | Laryngoscope set, neonate | 90 |  |
| 12 | Centrifuge, hematocrit, bench top, up to 4000 RPM, including rotor | 5 |  |
| 13 | Pump, suction, portable, 220V, w/access. | 46 |  |
| 14 | Stadio Meter | 29 |  |
| 15 | ABG Analysis Machine | 1 |  |
| 16 | Portable Ultra sound | 1 |  |
| 17 | Portable X - Ray machine | 1 |  |
| 18 | Flux meter | 1 |  |
| 19 | Nebulization Machine | 39 |  |
| 20 | CPAP | 5 |  |
| 21 | Fumigation Machine | 3 |  |
| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-03 |  |  |  |
| Blood Bank Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs: |
| 1 | Blood Bank Refrigerator (2 o C - 8 o C) | 9 | 20000.00 |
| 2 | Water Bath | 13 |  |
| 3 | Precision Incubator | 13 |  |
| 4 | Binocular Microscope | 11 |  |
| 5 | Table Top Centrifuge | 10 |  |
| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-04 |  |  |  |
| Diesel Generator Sets - 5KVA & 25 KVA |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs: |
| 1 | Generator set 25 KVA | 5 | 50000.00 |
| 2 | Generator 5 KVA | 25 |  |
| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-05 |  |  |  |
| Hospital Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs: |
| 1 | Infant meter, plexus, 3½ft/105cm | 4 | 10000.00 |
| 2 | Light, examination, mobile,220-12V | 2 |  |
| 3 | Spot Lamps | 9 |  |
| 4 | Crash Cart | 9 |  |
| 5 | Medicine Trolly | 10 |  |
| 6 | Basin, kidney, stainless steel, 825ml | 24 |  |
| 7 | Tray,dressing,ss,300x200x30mm | 23 |  |
| 8 | Glucometer with Dextrostix | 18 |  |
| 9 | X-Ray Viewing Box | 60 |  |
| 10 | Hub Cutter, syringe | 40 |  |
| 11 | E.Weighing Scale Digital - 100 Kg. | 29 |  |
| 12 | Scale, Baby, Electronic, 10 kg <5kg> | 67 |  |
| 13 | Weighing Scale Digital - 10Kg (Resolution up to 2 grams) | 2 |  |
| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-06 |  |  |  |
| General Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs: |
| 1 | Table | 71 | 10000.00 |
| 2 | Table & Chair | 3 |  |
| 3 | Refrigerator, hot zone, 210L | 6 |  |
| 4 | Wall Clock with second hand | 16 |  |
| 5 | Single Door Refrigerator (Domestic with Temp. Display) | 11 |  |
| 6 | Freezer | 12 |  |
| 7 | Ice Box | 12 |  |
| 8 | AC (1.5 Tonne) | 13 |  |

Sealed Tenders in will be accepted till 14.00 Hrs on 26/05/2014,

1. Interested eligible Bidders may obtain further information from the office of the Joint General Manager (Mtls), HLL Lifecare Ltd, Akkulam Factory, Thiruvananthapuram -17

2. Deleted.

3. The Tender Document can be downloaded from our Website www.lifecarehll.com.

4. Interested eligible bidders may submit their bidding documents to the office of the Joint General Manager (Mtls), HLL Lifecare Ltd, Akkulam Factory, Thiruvananthapuram –17 on the specified date and time. Bids will be evaluated for each item but tender-wise separately.

5. All bids must be accompanied by the items/documents specified herein and must be delivered to the office of Joint General Manager (Mtls) on or before 26/05/2014,      14.00 Hrs. The Technical Bids will be opened on 26/05/2014, at 14.30 Hrs, in the presence of the bidders or their duly authorized representatives who wish to attend the bid opening on the specified date and time at HLL's office mentioned above. In the event of the date being declared is a closed holiday for HLL, the due date for submission of bids and opening of bids will be the following working day at the appointed time.

SECTION II

INSTRUCTION TO BIDDERS

INSTRUCTION TO BIDDERS

A.INTRODUCTION

1. Eligible Bidders

  - Eligible to participate in the tender:

| Sl. No. | Minimum Eligibility criteria | Narrations Other than Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-06 |
| --- | --- | --- |
| 1 | Experience | Should have a minimum of 5 years experience in the field of dealing medical equipments. Should have the experience in supplying, installing and commissioning of high end hospital equipments. |
| 2 | Licenses/Registrations for undertaking, supply and installation contracts | - Primary manufactures who participate in the tender must produce their valid manufacturing license - Item wise Technical compliance sheet to match the technical specification prescribed in this tender. - Tenderers, other than manufacturers, who participate in this tender must produce the follows: - An authorization letter from Manufacturer’s / authorized distributors to participate in this tender. - Manufacturer’s manufacturing license / Distributor appointment letter to do the business in Himachal Pradesh State. |
| 3 | Quality Certifications | - Should have necessary valid Quality / System Certifications like ISO, CE, FDA etc) as applicable – to match technical specifications/requirements item wise as in the tender |
| 4 | Good financial capability necessary to perform the Contract | - Annual Turn over Statement preceding last three years (Certified by the Chartered accountant) - Profit & Loss Account for preceding last three years (Certified by the Chartered accountant) - Balance sheet for the past preceding last three years (Certified by the Chartered accountant) |

1.2 	Bidders should not be associated with the purchaser for the preparation 	of the design, specifications, and other documents used for the 	procurement of the goods under this Invitation of Bids.

1.3 	Deleted.

1.4 	Deleted.

2. Cost of Bidding

2.1 	The Bidder shall bear all costs associated with the preparation and submission of its bid and HLL Lifecare Limited, Akkulam Factory, Thiruvananthapuram hereinafter referred to as “the Purchaser”, will in no case be responsible or liable for these costs, regardless of the conduct or outcome of the bidding process.

B. THE BIDDING DOCUMENTS

3. Contents of Bidding Documents

3.1 	The goods required, bidding procedures and contract terms are prescribed in the Bidding documents. In addition to the Invitation for Bids, the Bidding documents include:

a.	Instruction to Bidders (ITB);

b.	General Conditions of Contract (GCC);

c.	Special Conditions of Contract (SCC);

d.	Technical Specifications;

e. 	Qualification criteria;

f.	Performance statement

g.	Bid Form;

h.	Price Schedule;

i.	Check List

3.2 	The Bidder is expected to examine all instructions, forms, terms and specifications in the Bidding Documents. Failure to furnish all information required by the Bidding Documents or submission of a bid not substantially responsive to the Bidding Documents in every respect will be at the Bidder’s risk and may result in rejection of its bid.

4. Clarification of Bidding Documents

4.1  	A prospective Bidder requiring any clarification of the Bidding Documents may notify the Purchaser in writing at the Purchaser’s mailing address indicated in the Invitation for Bids. The Purchaser will respond in writing to any request for clarification of the Bidding Documents which it receives not later than 5 days prior to the deadline for submission of bids prescribed by the Purchaser. Any such response (relevant) against the clarification of the supplier shall be notified in our website only.

5. Amendment of Bidding Documents

5.1  	At any time prior to the deadline for submission of bids, the Purchaser may, for any reason, whether at its own initiative or in response to a clarification requested by a prospective bidder, modify the Bidding Documents by an amendment.

Those bidders who download the tender documents from our website should furnish the Name and address of the vendor, name of contact person, telephone numbers and Email details immediately by Email to HLL Akkulam. Any changes pertains to this tender shall be communicated only through our website www.lifecarehll.com

5.2 	All prospective bidders who have received the Bidding Documents will be notified of the amendment in writing or by mail and will be binding on them.

5.3 	In order to allow prospective bidders reasonable time in which to take the amendment into account in preparing their bid, the Purchaser may, at its discretion, extend the deadline for the submission of bids

C. PREPARATION OF BIDS

6. Language of Bid

6.1 	The Bid prepared by the bidder and all correspondence and documents relating to the bid exchanged by the Bidder and the Purchaser, shall be written in the English language. Supporting documents and printed literature furnished by the Bidder may be written in another language provided they are accompanied by an accurate translation of the relevant passages in the English language in which case, for purposes of interpretation of the Bid, the English translation shall govern

7. Documents Comprising the Bid

Following documents and forms are to be included in the Technical bid and Commercial bid respectively.

7.1 TECHNICAL BID

The technical bid shall consist of the following documents.

- Tender Fee : DD for Rs. 5000.00 (inclusive of tax) drawn in favour of HLL Lifecare Limited payable at Thiruvananthapuram is to be attached towards tender fee.

- EMD / Bid Security in form of DD, drawn in favor of HLL Lifecare Limited, payable a Thiruvananthapuram. The EMD / Bid Security shall be refunded to the non-responsive bidders within 60 days from the date of opening of Bid. The Tender received without EMD will be rejected.

- The technical specifications enclosed along with the bid document shall be confirmed by signature of the bidder/authorized signatory of the bidding firm, in all pages and authorized by official seal. The information shall be filled, wherever necessary as per the instructions given.

- Duly attested copies of factory license/Industrial/licenses, sales tax registration, and documents to prove the legal status, place of registration and principal place of business of the undertaking

- Duly attested copies of quality certificates for the products, quality system certifications as specified in technical specification. Item wise Product catalogue/Detailed printed literature to prove the technical specification along with COMPLIANCE SHEET confirming to technical specification should be attached with offer.

- Copy of Balance sheet for the past three financial years, duly certified by a chartered accountant

- Performa for equipment and quality control, duly filled as per section VII

- Performance statement in the proforma under section VIII

- Documentary evidence established in accordance with ITB Clause 1 that the Bidder is eligible to bid and is qualified to perform the contract if its bid is accepted;

- Documentary evidence established in accordance with ITB Clause 12 that the goods and ancillary services to be supplied by the Bidder shall conform to the Bidding Documents

- Duly filled Check list as per section XII

7.2  PRICE BID

Duly filled Bid Form and price schedule as per sections IX, X and XI respectively.

The Tenderer shall indicate on the Price Schedule provided under Section X all the specified components of prices shown therein including the unit prices and total tender prices of the goods and services it proposes to supply against the requirement. All the columns shown in the price schedule should be filled up as required. If any column does not apply to a tenderer, same should be clarified as “NA” by the tenderer.

The bidders have to quote for all the equipments in the respective tender, while quoting. Say, if the bidder is quoting for Tender for Hospital Equipments, the bidder has to quote for all equipments in the particular tender.

- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-01- Neonate Critical Equipments

- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-02 - Newborn Unit Equipments

- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-03 - Blood Bank Equipments

- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-04 - Diesel Generator Sets - 5KVA & 25 KVA

- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-05- Hospital Equipments

- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-06 - General Equipments

The purchaser will have the right to award the contract as per the eligibility and to the best benefit of the exchequer.

The Price schedule for all six tenders shall be kept in separate sealed envelopes is mandatory.

The tenderer has the option to submit its quotation for any one or more Tender. However, while quoting for a Tender, the tenderer shall quote for the complete requirement of goods and services as specified in that particular Tender. All sundry equipments, fittings, units assemblies, accessories and all other items which are useful and necessary for efficient assembly and installation of equipment and components of the work be deemed to have been included in the tender irrespective of the fact whether such items are specifically mentioned in the tender documents or not.

While filling up the columns of the Price Schedule, the following aspects should be noted for compliance:

a) any sales or other taxes and any duties including excise duty, which will be payable on the goods if the contract is awarded;

b) charges towards Packing & Forwarding, Inland Transportation, Insurance (local transportation and storage) would be borne by the Contractor from ware house to the project site including Loading/Unloading and other local costs incidental to delivery of the goods to the destination.

c) the price of Incidental Services, as mentioned in requirements and Price Schedule;

e) the prices of Turnkey ( if any), a s mentioned in requirements, Technical Specification and Price Schedule;

f) The rates quoted by the tenderer, shall be firm and fixed and inclusive of all taxes including work contract taxes, duties and levies and all charges for packing forwarding, insurance , freight and delivery , installation , testing commissioning at site and clearance from local authorities etc.

g) the price of annual CMC, as mentioned in requirements, Technical Specification and Price Schedule.

8. Bid Form

8.1 	The Bidder shall fill up the Bid Form and the appropriate Price Schedule furnished in the bidding documents, indicating, a brief description of the equipment, its country of origin etc.

9.  Bid Prices

9.1    The Bidder shall indicate in the Price Schedule the unit prices and total Bid prices of the equipment it proposes to supply under the Contract.

10. Bid Currencies

10.1 Prices shall be quoted in Indian Rupees.

11. Documents establishing bidder’s eligibility and qualifications

11.1 	Pursuant to ITB Clause 7, the bidder shall furnish, as part of its bid, documents establishing the bidder’s eligibility to bid and its qualifications to perform the Contract if its bid is accepted.

11.2 	The documentary evidence of the Bidder’s qualifications to perform the Contract if its bid is accepted, shall establish to the HLL’s satisfactions.

(a) that the bidder has the financial, technical and production capability necessary to perform the Contract and meets the criteria outlined in the qualification requirements specified in Section VI. To this end, all bids submitted shall include the following information.

(i) The legal status, place of registration and principle place of business of the company or firm or partnership, etc;

(ii) Details of experience and past performance of the bidder on items offered and on those of similar nature within the past Five years and details of current contracts in hand and other commitments (as per proforma given in Section VIII)

12 .Documents establishing equipment’s conformity to bidding documents

12.1 	Pursuant to ITB Clause 7, the Bidder shall furnish, documents establishing the conformity to the bidding documents of all equipment and services which the bidder proposes to supply under the contract.

12.2   Deleted

12.3  Deleted

13. Period of Validity of Bids

13.1  Price Bid shall be valid for One year from the date of opening of Price Bid

13.2 	In exceptional circumstances, the Purchaser may solicit the bidders consent to an extension of the period of validity. The request and the responses thereto shall be made in writing.

14. Format and Signing of Bid

14.1 	The bidder shall prepare the bid clearly marking each “Technical Bid” and “Price Bid” as appropriate.

14.2   The bid shall be typed or written in indelible ink and shall be signed by the Bidder or a person or persons duly authorized to bind the Bidder to the Contract. The letter of authorization shall be indicated by written power-of-attorney accompanying the bid. All pages of the bid, except for unamended printed literature, shall be initialed by the person or persons signing the bid.

14.3   Any interlineations, erasures or overwriting shall be valid only if they are initialed by the person or persons signing the bid.

D. SUBMISSION OF BIDS

#### 15. Sealing and Marking of bids

Separate bids shall be submitted for price bid and technical bids.

15.1   The bidders shall seal Technical bid   and Price bid duly marking the envelopes as “TECHNICAL BID” - cover “A” and “PRICE BID” - cover “B”.  (Tender wise separate sealed envelopeshas to be placed in a single cover – as price bid cover  B ”. The Bidders shall then place both the envelopes (cover “A” & cover “B”) in an outer envelope. The name of the product for which the bid is made must be marked on both the inner envelops and the outer envelope.

15.2   The inner and outer envelopes shall be addressed to the Purchaser at the following address:

Joint General Manager (Materials)

HLL Lifecare Limited, Akkulam Factory, Sreekaryam  P.O. ,

Thiruvananthapuram -17   Phone 0471 2445930 / 2445935

Email :tradingpaft@lifecarehll.com

- The outer envelope shall bear the Invitation for bids (IFB) number, and a statement: “DO NOT OPEN BEFORE” 14.30 Hrs. ON 26/05/2014   ”

- The inner envelope shall also indicate the name and address of the bidder to enable the bid to be returned unopened in case it is declared “late/disqualified”.

- If the inner and outer envelopes are not sealed and marked as required, the Purchaser will assume no responsibility for the bid’s misplacement or premature opening.

- Bids must be received by the Purchaser at the address specified not later than the date and time specified in the Invitation For Bid.  In the event of the specified date for the submission of bids, being declared a holiday for the Purchaser, the bids will be received up to the appointed time on the next working day.

- The Purchaser may, at its discretion, extend this deadline for the submission of bids by amending the bidding documents in which case all rights and obligations of the Purchaser and bidders previously subject to the deadline will thereafter be subject to the deadline as extended.

16. LATE BIDS

16.1   Any bid received by the Purchaser after the deadline for submission of bids prescribed by the Purchaser, will be rejected and returned unopened to the bidder.

17.  MODIFICATION AND WITHDRAWAL OF BIDS

17.1 	The bidder may modify or withdraw its bid after the bid submission, provided that written notice of the modification or withdrawal is received by the Purchaser prior to the deadline prescribed for submission of bids.

17.2   The bidder’s modification or withdrawal notice shall be prepared, sealed, marked and dispatched in accordance with the provisions of ITB Clause 15.  A withdrawal notice may also be sent by email, but followed by a signed confirmation copy, post-marked not later than the deadline for submission of bids.

17.3   No bid may be modified subsequent to the deadline for submission of bids.

17.4 	No bid may be withdrawn in the interval between the deadline for submission of bids and the expiration of the period of bid validity specified by the bidder in the Bid Document.

E. BID OPENING AND EVALUATION

18. Opening of Bids by Purchaser

18.1 The Purchaser will open all technical bids (Cover A), in the presence of bidder's representatives who choose to attend, at 14.30 Hrs on 26/05/2014    at the following location:

Joint General Manager (Materials)

HLL Lifecare Limited, Akkulam Factory, Sreekaryam  P.O. ,

Thiruvananthapuram -17  Phone 0471 2445930 / 2445935

Email : tradingpaft@lifecarehll.com

18.2 	The bidder’s representatives who are present shall sign a register evidencing their attendance. In the event of the specified date of bid opening being declared a holiday for the Purchaser, the bids shall be opened at the appointed time and location on the next working day.

18.3   The bidder’s names, modifications, bid withdrawals and such other details as the Purchaser, at its discretion, may consider appropriate will be announced at the opening. No bid shall be rejected at bid opening, except for late bids, which shall be returned unopened to the bidder pursuant to ITB Clause 16.

18.4    Deleted

18.5   The “Price Bid” (Cover B) will be opened after evaluation of “Technical bids” (Cover A) and the date and time will be intimated to bidders whose bids are qualified and responsive.

18.5 	The purchaser will scrutinize the technical bid for compliance to the specifications and documentation requirement as per the bid document. Purchaser will depute its competent officers to the premises of the bidder qualified on the basis of technical scrutiny, for on-site evaluation of the claims made in the technical bid, if deemed appropriate on purchaser’s sole discretion. The bidders will be short-listed on the basis of responsiveness of technical bid as well as report of on-site technical evaluation. The price bid of the bidders who are disqualified at the technical scrutiny and on-site evaluation will be returned un-opened. The on-site evaluation may include the inspection of the specimen sample of the equipment. The short listed bidders will be informed about the time, date and venue of the price bid opening. The successful bidder shall be identified on the basis of lowest evaluated substantially responsive bid.

19.  CLARIFICATION OF BIDS

19.1   During evaluation of bids, the Purchaser may, at its discretion, ask the bidder for a clarification of its bid.  The request for clarification and the response shall be in writing and no change in prices or substance of the bid shall be sought, offered or permitted.

20   PRELIMINARY EXAMINATION

  - The Purchaser will examine the bids to determine whether they are complete, whether any computational errors have been made, whether the documents have been properly signed, and whether the bids are generally in order.

20.2   Arithmetical errors will be rectified on the following basis.  If there is a discrepancy between the unit price and the total price that is obtained by multiplying the unit price and quantity, the unit price shall prevail and the total price shall be corrected. If there is a discrepancy between words and figures, the amount in words will prevail.  If the bidder does not accept the correction of the errors, its bid will be rejected.

20.3 	The Purchaser may waive any minor informality or non-conformity or irregularity in a bid, which does not constitute a material deviation, provided such a waiver does not prejudice or affect the relative ranking of any bidder.

20.4   	Prior to the detailed evaluation, pursuant to Clause ITB 21, the purchaser will determine the substantial responsiveness of each bid to the bidding documents. For purposes of these clauses a substantially responsive bid is one which conforms to all the terms and conditions of the bidding documents without material deviations. Deviations from or objections or reservations to critical provisions such as those concerning Warranty, Force Majeure, Applicable law  and Taxes and Duties  will be deemed to be material deviation. The purchaser’s determination of a bid’s responsiveness is to be based on the contents of the bid itself without recourse to extrinsic evidence.

20.5   If a bid is not substantially responsive, it will be rejected by the Purchaser and may not subsequently be made responsive by the bidder by correction of the non-conformity.

21. EVALUATION AND COMPARISON OF BIDS

21.1   The Purchaser will evaluate and compare bids previously determined to be substantially responsive, pursuant to ITB Clause 20.

21.2   The purchaser’s evaluation of a bid will take into account the bid price including all costs as well as duties and taxes paid or payable and price of incidental services, if any, in the manner and to the extent indicated in ITB Clause 21.3 and in the technical specifications:

i.  cost of transportation, insurance and other costs incidental to the delivery of goods; to designated delivery destination and installation and handing over the same to the Purchaser, at designated Hospitals in Himachal Pradesh State.

ii. Deleted

21.3 	Pursuant to ITB Clause 21.2 the following evaluation methods will be applied:

(a) transportation, from factory, insurance and   incidentals.

(i) transportation, insurance and other incidentals, for delivery of goods to the required locations mentioned in ITB Section 21.2 The above costs will also be considered in the bid price.

(b)  Spare parts and after sales service facilities:

The cost of the Purchaser of establishing the minimum service facilities and parts inventories, as outlined elsewhere in the bid invitation, if quoted separately, shall be added to the bid price.

(c) Comprehensive Maintenance Contract (CMC)

The Purchaser desires to have separately comprehensive maintenance charges for a period of 5 years after the expiry of free maintenance period, clearly indicating year wise comprehensive maintenance charges, as per Section XI.

22. CONTACTING THE PURCHASER

22.1     No bidder shall contact the Purchaser on any matter relating to its bid, from the time of the bid opening to the time the contract is awarded.    If the bidder wishes to bring additional information to the notice of the purchaser it should do so in writing.

22.2    Any effort by a bidder to influence the purchaser in its decisions on bid evaluation, bid comparison, or selection may result in the rejection of the bidders bid.

F. AWARD OF CONTRACT

23. POST QUALIFICATION

23.1 	In the absence of pre-qualification, the Purchaser will determine to its satisfaction whether the bidder that is selected as having submitted the lowest evaluated responsive bid meets the criteria specified in ITB Clause 11.2 and is qualified to perform the contract satisfactorily

23.2 	The determination will take into account the bidders financial, technical and production/other capabilities. It will be based upon an examination of the documentary evidence of the bidder’s qualifications submitted by the bidder, pursuant to ITB Clause 11 as well as such other information as the Purchaser deems necessary and appropriate.

23.3 	An affirmative determination will be a prerequisite for award of the contract to the bidder. A negative determination will result in rejection of the bidders bid, in which event the Purchaser will proceed to the next lowest evaluated bid to make a similar determination of that bidder’s capabilities to perform satisfactorily.

24. AWARD CRITERIA

24.1 	Subject to ITB Clause 28, the Purchaser will award the contract to the successful bidder whose bid has been determined to be substantially responsive and has been determined as the lowest evaluated bid, provided further that the bidder is determined to be qualified to perform the contract satisfactorily.

25. PURCHASER’S RIGHT TO VARY QUANTITIES AT TIME OF AWARD

  - The Purchaser reserves the right at the time of award of contract to increase or decrease the quantity of goods and services originally specified in the bid document without any change in unit price or other terms and conditions with in the bid validity period.

  - The quantity mentioned herein are approximate present requirement and in case more or less quantity is required, the suppliers should be prepared to effect supply at short notice on the same terms and conditions.

26. PURCHASER’S RIGHT TO ACCEPT ANY BID AND TO REJECT ANY OR ALL BIDS

26.1 		The Purchaser reserves the right to accept or reject any bid, and to annul the bidding process and reject all bids at any time prior to award of Contract, without thereby incurring any liability to the affected bidder or bidders.

27. NOTIFICATION OF AWARD

27.1 	Prior to the expiration of the period of bid validity, the Purchaser will notify the successful bidder in writing by registered letter or by cable, to be confirmed, that its bid had been accepted.

27.2	The notification of award will constitute the formation of the contract

28. SIGNING OF SUPPLY AGREEMENT

28.1 	The successful bidder has to sign a supply agreement, the draft of which will be send to the successful bidder only with the letter informing the acceptance of bid. Within 7 days of receipt of the information regarding acceptance of the bid, the successful bidder shall sign the supply agreement and return it to the Purchaser.

29. DELAYS IN EXECUTION OF ORDER

29.1	The supplier shall supply, install and hand over the entire equipment on turnkey basis  within 30 days from the date of placing the Purchase Order/LOI.

29.2 	A delay by the Supplier in the performance of its delivery obligations shall render the Supplier liable to the imposition of penalty pursuant to agreement, unless an extension of time is agreed upon pursuant to agreement without the application of liquidated damages.

  - If the Supplier fails to deliver any or all of the equipment or perform the Services within the time period(s) specified in the Purchase Order, the purchaser shall, without prejudice to its other remedies under the Contract, deduct from the Contract Price, as penalty, a sum equivalent to the percentage specified in the SCC clause 10.

SECTION III

GENERAL CONDITIONS OF CONTRACT

GENERAL CONDITIONS OF CONTRACT

1. DEFINITIONS

1.1	In this contract the following terms shall be interpreted as indicated:

(a)  “The Contract” means the agreement entered into between the Purchaser and the Supplier as recorded in the Contract Form signed by the parties, including all the attachments and appendices thereto and all documents incorporated by reference therein;

(b) “The Contract Price” means the price payable to the Supplier under the Contract for the full and proper performance of its contractual obligations;

(c) “The Goods” means all the products, equipment, machinery, and/or other materials which the Supplier is required to supply to the Purchaser under the Contract;

(d) “Services” means services ancillary to the supply of the Goods, such as transportation and insurance, and other incidental services, such as installation, commissioning, provision of technical assistance, training and other obligations of the Supplier covered under the contract;

(e) “GCC” means the General Conditions of Contract contained in this section.

(f) “SCC” means the Special Conditions of Contract.

(g) “The Purchaser” means the Organization purchasing the Goods, as named in SCC;

(h) “The Supplier” means the individual or firm supplying the Goods under this Contract;

(i)  “Day” means calendar day.

(j)  “Delivery period” means the period applicable upto completion of supply; installation and testing by the supplier at the required location mentioned in purchase order and accepted by the Purchaser.

(k) “EarnestMoneyDeposit”(EMD) means Bid Security /monetary or financial guarantee to be furnished by a tenderer along with Technical Bid Documents.

2. APPLICATION

2.1	These General Conditions shall apply to the extent that they are not superseded by provisions in other parts of the Contract.

3. STANDARDS

3.1	The equipment supplied under this Contract shall conform to the standards mentioned in the Technical Specifications, and, when no applicable standard is mentioned, to the authoritative standard appropriate to the Goods’ country of origin and such standards shall be the latest issued by the concerned institution.

4. USE OF CONTRACT DOCUMENTS AND INFORMATION

4.1 	The Supplier shall not, without the Purchaser’s prior written consent, disclose the Contract, or any provision thereof, or any specification, plan, drawing, pattern, sample or information furnished by or on behalf of the Purchaser in connection therewith, to any person other than a person employed by the Supplier in performance of the Contract. Disclosure to any such employed person shall be made in confidence and shall extend only so far as may be necessary for purposes of such performance.

4.2	The Supplier shall not, without the Purchaser’s prior written consent, make use of any document or information enumerated in GCC Clause 4.1 except for purposes of performing the Contract.

4.3	Any document, other than the Contract itself, enumerated in GCC clause 4.1 shall remain the property of the Purchaser and shall be returned (in all copies) to the Purchaser on completion of the supplier’s performance under the Contract if so required by the Purchaser.

5. PATENT RIGHTS

5.1	The Supplier shall indemnify the Purchaser against all third-party claims of infringement of patent, trademark or industrial design rights arising from use of the Goods or any part thereof in India.

6. INSPECTION AND TESTS

6.1	The Purchaser or its representative shall have the right to inspect and/or to test the Goods to confirm their conformity to the Contract at no extra cost of the Purchaser. The Special conditions of Contract and/or the Technical Specifications shall specify what inspections and tests the Purchaser requires and where they are to be conducted. The Purchaser shall notify the Supplier in writing of the identity of any representatives retained for these purposes.

6.2	The inspections and test may be conducted on the premises of the Supplier or its subcontractor(s), at point of delivery and/or at the Goods final destination. Where conducted on the premises of the Supplier or its subcontractor(s), all reasonable facilities and assistance including access to drawings and production data - shall be furnished to the inspectors at no charge to the Purchaser.

6.3	Should any inspected or tested Goods fail to conform to the specifications, the Purchaser may reject them and the Supplier shall either replace the rejected Goods or make all alternations necessary to meet specification requirements free of cost to the Purchaser.

6.4 	The Purchasers right to inspect, test and, where necessary, reject the Goods’ arrival in at site shall in no way be limited or waived by reason of the Goods having previously been inspected, tested and passed by the Purchaser or its representative prior to the Goods dispatched.

7. PACKING

7.1 	Deleted.

7.2 	Deleted

8 DELIVERY AND DOCUMENTS

8.1	Delivery destination (for present estimated requirement) of the Goods is with in designated Hospitals in Himachal Pradesh State (Designated location with item wise per unit quantity - [as detailed in Section XVIII Consignee details]. The equipment supplied and installed shall be in accordance with the terms specified in the purchase order by the Purchaser. The documents like warranty certificate, equipment manual, test certificate and also other documents to be furnished by the Contractor as specified in the purchase order.

9. INSURANCE

9.1 	Deleted

10.  INCIDENTAL SERVICES

10.1 	The supplier may be required to provide any or all of the following services, including additional services, if any, specified in SCC:

(a) Performance or supervision of the on-site assembly and/or start-up of the supplied goods

(b) Furnishing of tools required for assembly and/or maintenance of the supplied Goods;

(c) Furnishing of detailed operations and maintenance manual for each product

(d) performance or supervision or maintenance and/or repair of the supplied Goods, for a period of time agreed by the parties, provided that this service shall not relieve the Supplier of any warranty obligations under this Contract; and

(e) Training of the Purchaser’s Personnel, at the Supplier’s plant and/or on-site, in assembly, start-up, operation, maintenance and/or repair of the supplied Goods.

11 SPARE PARTS

11.1 	As specified in the Special Conditions of Contract, the Supplier may be required to provide any or all of the following materials and notifications and information pertaining to spare parts manufactured or distributed by the Supplier:

(a)Such spare parts as the Purchaser may elect to purchase from the Supplier, providing that this election shall not relieve the Supplier of any warranty obligations under the Contract; and

(b)In the event of termination of production of the spare parts:

(i) Advance notification to the Purchaser of the pending termination, in sufficient time to permit the Purchaser to procure needed requirements; and

(ii) Following such termination, furnishing at no cost to the Purchaser,  the blueprints, drawings and specifications of the spare parts, if and when requested.

12. WARRANTY

12.1 	The Supplier warrants that the Goods supplied under this Contract are new, unused, of the most recent or current models and incorporate all recent improvements in design and materials unless provided otherwise in the Contract. The supplier further warrants that the Goods supplied under this Contract shall have no defect arising from design, materials or workmanship (except when the design and/or material is required by the Purchaser’s specifications) or from any act or omission of the Supplier, that may develop under normal use of the supplied Goods in conditions obtaining in the state/country of final destination

12.2 	This warranty shall remain valid for one year after the Goods or any portion thereof as the case may be, have been delivered to and accepted at the final destination indicated in the Contract.

12.3 	The Purchaser shall promptly notify the supplier in writing of any claims arising under this warranty.

12.4 	Upon receipt of such notice, the Supplier shall, within the period specified in SCC and with all reasonable speed, repair or replace the defective Goods or parts thereof, without cost to the Purchaser

12.5 	If the Supplier, having been notified, fails to remedy the defect(s) within the period specified in SCC within a reasonable period, the Purchaser may proceed to take such remedial action as may be necessary, at the Supplier’s risk and expense and without prejudice to any other rights which the Purchaser may have against the Supplier under the contract.

  - Deleted

13. PAYMENT

13.1	The method and conditions of payment to be made to the Supplier under the Contract shall be specified in the Special Conditions of Contract.

13.2	The Suppliers request(s) for payment shall be made to the Purchaser in writing, accompanied by an invoice describing, as appropriate, the Goods delivered and the services performed, and by documents, submitted pursuant to GCC Clause 8, and upon fulfillment of other obligations stipulated in the contract.

13.3	Payment shall be made promptly by the Purchaser but in no case later than sixty (60) days after submission of the invoice/claim by the Supplier.

13.4 	Payment shall be made in Indian Rupees

14. PRICES

14.1	Prices charged by the Supplier for Goods delivered and Services performed under the Contract shall not vary from the prices quoted by the Supplier in its bid, with the exception of any prices adjustments authorized in the special Conditions of Contract or in the Purchaser’s request for bid validity extensions, as the case may be.

15. CHANGE ORDERS

15.1	The Purchaser may at any time by written order given to the Supplier, make changes within the general scope of the Contract in any one or more of the following:

- drawings, designs or specifications, where Goods to be furnished under the Contract are to be specifically executed for the Purchaser;

- deleted

- the place of delivery; or

- the services to be provided by the Supplier.

15.2 	If any such changes cause an increase or decrease in the cost of, or the time required for, the Supplier’s performance of any provisions under the Contract, an equitable adjustment shall be made in the Contract Price or delivery schedule, or both, and the Contract shall accordingly be amended. Any claims by the Supplier for adjustment under this clause must be asserted within thirty (30) days from the date of the Supplier’s receipt of the Purchaser’s change order.

16. CONTRACT AMENDMENTS

16.1 	Subject to GCC Clause 17, no variation in or modification of the terms of the Contract shall be made except by written amendment signed by the parties.

17. ASSIGNMENT

17.1 	The Supplier shall not assign, in whole or in part, its obligations to perform under the contract, except with the Purchaser’s prior written consent.

18. SUBCONTRACTS

18.1	The supplier shall notify the Purchaser in writing of all subcontracts awarded under the contract if not already specified in his bid. Such notification, in his original bid or later, shall not relieve the Supplier from any liability or obligation under the contract.

19. LIQUIDATED DAMAGES

19.1	Subject to SCC clause 10.1..

20. TERMINATION BY DEFAULT

20.1	The Purchaser may, without prejudice to any other remedy for breach of contract, by written notice of default sent to the supplier, terminate the Contract in whole or part;

(a)if the Supplier fails to execute any or all of the contracts within the time period(s) specified in the Contract, or within any extension thereof granted by the Purchaser, or

(b) if the Supplier fails to perform any other obligation(s) under the contract.

20.2	In the event the Purchaser terminates the Contract in whole or in part,  the Purchaser may procure, upon such terms and in such manner as it deems appropriate, Goods or Services similar to those unexecuted, and the Supplier shall be liable to the Purchaser for any excess costs for such similar Goods. However, the Supplier shall continue the performance of the Contract to the extent not terminated.

21.  FORCE MAJEURE

21.1	For purposes of this Clause “Force Majeure” means an event beyond the control of the Supplier and not involving the Supplier’s fault or negligence and not foreseeable. Such events may include, but are not limited to, acts of the Purchaser either in its sovereign or contractual capacity, wars or revolutions, fires, floods, epidemics, quarantine restrictions and freight embargoes.

21.2	If a Force Majeure situation arises, the Supplier shall promptly notify the Purchaser in writing of such conditions and the cause thereof. Unless otherwise directed by the Purchaser in writing, the Supplier shall continue to perform its obligations under the Contract as far as is reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the force majeure event.

22. TERMINATION FOR INSOLVENCY

22.1 	The Purchaser may at any time terminate the Contract by giving written notice to the Supplier, if the Supplier becomes bankrupt or otherwise insolvent. In this event, termination will be without compensation to the Supplier, provided that such termination will not prejudice or affect any right of action or remedy, which has accrued or will accrue thereafter to the Purchaser.

23. RESOLUTION OF DISPUTES

23.1 	The Purchaser and the supplier shall make every effort to resolve amicably by direct informal negotiation any disagreement or dispute arising between them under or in connection with the Contract.

  - If, after thirty (30) days from the commencement of such informal negotiations, the Purchaser and the Supplier have been unable to resolve amicably a Contract dispute, either party may require that the dispute be referred for resolution to the formal mechanisms specified in the Special Conditions of Contract. These mechanisms may include, but or not limited to, conciliation mediated by a third Party, adjudication in an agreed national forum, and national arbitration.

  - The jurisdiction of any disputes, suits and proceeding arising out of the tender shall be only in the court of Thiruvananthapuram.

24. GOVERNING LANGUAGE

24.1 	Subject to ITB clause 6.1

25. APPLICABLE LAW

25.1 	The Contract shall be interpreted in accordance with the laws of the Union of India.

26. NOTICES

26.1 	Any notice given by one party to the other pursuant to this Contract shall be sent to other party in writing or by email and confirmed in writing to the other Party’s address specified in Special Conditions of Contract.

26.2 	A notice shall be effective when delivered or on the notice’s effective date, whichever is later.

27. TAXES AND DUTIES

27.1	Suppliers shall be entirely responsible for all taxes, duties, license fees, octroi etc., incurred until delivery of the contracted Goods to the Purchaser.

SECTION IV

SPECIAL CONDITIONS OF CONTRACT

SPECIAL CONDITIONS OF CONTRACT

The following Special Conditions of Contract shall supplement the General Conditions of Contract. Whenever there is a conflict, the provisions herein shall prevail over those in the General Conditions of contract. The Corresponding clause number of the General Conditions is indicated in parentheses.

1. DEFINITIONS (GCC Clause 1)

(a) The Purchaser is HLL Lifecare Limited, Akkulalm Factory, Thiruvananthapuram

2. INSPECTION AND TESTS (GCC Clause 6)

2.1	The following inspection procedures and tests are required by the Purchaser; the supplier shall get each equipment inspected in manufacturer’s works and submit a test certificate and also guarantee/warranty certificate that the equipment conforms to laid down specifications.

2.2	The Purchaser or its representative shall inspect and/or test any or all the equipment to confirm their conformity to the Contract specifications, prior to dispatch from the manufacturer’s premises. Such inspection and

clearance will not prejudice the right of the consignee to inspect and test the equipment on receipt at destination.

2.3	If the equipment fails to meet the laid down specifications the supplier shall take immediate steps to remedy the deficiency or replace the defective equipment to the satisfaction of the Purchaser.

3. INSURANCE (GCC Clause 9)

3.1	Deleted

4. INCIDENTAL SERVICE (GCC Clause 10)

4.1	The following services covered under GCC Clause 10 shall be furnished and the cost shall be included in the contract price:

(a)Unloading, safe storage and handling of consignment off site.

(b)On site assembly of the supplied goods, installation, testing and

commissioning of the equipment.

(c)Furnishing of detailed operations and maintenance manual for each appropriate unit of supplied Goods;

5. SPARE PARTS (GCC Clause 11)

Add as Clause 11.2 to the GCC the following:

5.1	Supplier shall carry sufficient inventories to assure ex-stock supply of consumable spares for the Goods. Other main spare parts and components shall be supplied as promptly as possible but in any case within one week of placement of order.

6. WARRANTY (GCC Clause 12)

Substitute GCC Clause 12.2 by the following:-

6.1	This warranty shall remain valid for ONE year after goods or any portion thereof as the case may be, have been delivered to and accepted at the final destination indicated in the contract.

6.2      Any major repair pointed out by the Purchaser within the warranty period shall be rectified by the Supplier from the date of intimation within a period of 3 calendar days and commission the equipment to the satisfaction of the Purchaser. Failing which the Purchaser has the right to levy penalty on the Supplier per day or part thereof for each equipment until the equipments are repaired and commissioned to the satisfaction of the Purchaser.

6.3	The Supplier shall, in addition, comply with the performance and/ or consumption guarantees specified under the contract. If for reasons attributable to the Supplier, these guarantees are not attained in whole or in part, the Supplier shall at its discretion either:

- make such changes, modifications, and/or additions to the Goods or any part thereof as may be necessary in order to attain the contractual guarantees specified in the Contract at its own cost and expense and to carry out further performance tests in accordance with SCC 2:

  - Deleted.

7. PAYMENT (GCC Clause 13)

7.1 	Payment for Goods and Services shall be made in Indian Rupees as follows:

- No advance payment is payable.

- 90 % payment shall be made within 30 days on supply of equipments in good and acceptable condition at the respective warehouse/Hospitals at designated delivery destinations. The balance 10 % of the invoice value shall be paid against submission of necessary documents  and Performance Bank Guarantee for equal value, valid for one year from the date of completion and handing over of equipments on turnkey basis to the Purchaser, as applicable. Total payment (100% payment) will be released only after satisfactory completion of Installation, commissioning, and handing over/final acceptance certificate from the the concerned hospital authorities.

8. PRICES (GCC Clause 14)

Substitute Clause 14.1 of the GCC with the following:

8.1	Prices payable to the Supplier as stated in the Contract shall not be subject to adjustment during performance of the Contract.

9. SUB CONTRACT (GCC Clause 18)

Add at the end of sub-clause 18.1 the following:

9.1	Sub-contract shall be only for bought-out items and sub-assemblies.

10. LIQUIDATED DAMAGES (GCC Clause 19)

For delays:

Substitute GCC Clause 19.1 by the following:

10.1	If the Supplier fails to deliver any or all of the Goods or perform of services within the time period(s) specified in the Contract, the Purchaser shall without prejudice to its other remedies under the Contract, deduct from the Contract Price, as liquidated damages, a sum equivalent to 0.5 percent of the delivered price of the delayed Goods or unperformed Services for each week of delay or part thereof until actual delivery or performance, up to a maximum deduction of 10 percent of the delayed Goods or Services contract price. Once the maximum is reached, the Purchaser may consider termination of the Contract.

11. Deleted.

12. COMPREHENSIVE MAINTENANCE CONTRACT (CMC):

12.1	During CMC period, any major repair pointed out by the Purchaser shall be rectified by the Supplier from the date of intimation within a period of 3 calendar days and commission the equipment to the satisfaction of the Purchaser. Failing which the Purchaser has the right to levy penalty on the Supplier per day or part thereof for each equipment until the equipments are repaired and commissioned to the satisfaction of the Purchaser.

12.2	The Supplier shall indicate clearly the free guarantee maintenance of the whole system supplied by the Supplier and the same should not be less than 1 year.

12.3	The Supplier shall also indicate separately post guarantee maintenance cost (CMC) of the entire system for 5 years subsequent to free maintenance period and shall clearly indicate year wise maintenance cost with probable cost of spares required for each year, in addition to comprehensive maintenance charges, as per Section XI.

SECTION V

TECHNICAL SPECIFICATIONS

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-01 |  |  |
| --- | --- | --- |
| Neonate Critical Equipments |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty |
| 1 | Open care system: Radiant warmer, fixed height, with trolley, drawers, stand infusion double hook two poles attached. | 85 |
| 2 | Phototherapy unit, Double Surface, tachable , LED Type | 12 |

1). Open care warmer system: Radiant warmer, with fixed height, trolley, drawers, 2 No’s side tray for mounting accessories, 2 IV pole stand with 4 hooks attached.

Technical Specifications:

- Mobile newborn resuscitation table with fixed-height radiant warmer

- Antistatic castors, 2 with breaks

- Table surface with mattress with infant head/shoulder support

- Mattress-padding: foam density approx. 21 - 25 kg/m3

- Mattress cover: removable with zipper, waterproof, washable, resistant to cleaning with    chlorine  based solution and flame retardant

- Side boards transparent acryl, drop down and lockable

- Under table 2 storage drawers

- 2 side trays for mounting of accessories like pulse oxymeter etc.

- Hood suspended above the table integrates heating element and overhead light

- Overhead light: 2 x 50W halogen spot, with dimming function

- Heating element: emitter with parabolic reflector and protected by metal grid

- Control unit allows air and skin temperature preset (LED indicator) and drives radiant heater output (servo and manual)

- Integrated timer: 1 to 59 min, with count-up and count-down feature

- Temperature range, skin: 34°C to 38°C (user pre-settable)

- Monitoring of skin temperature by means of sensor, range: 30°C to 42°C

- Heater output: 0 to 100% in increments of 5%

- Control unit: audiovisual alarms according to timer and temperature presets avoiding overheating

- Display reports systems errors, temperature sensor failure

- Power requirement: 220 V/50 Hz

- Power consumption: 800 W

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE, FDA 510k or equivalent

Supplied with:

- 1 x skin temperature probe (including connection cable)

- 1 x spare heating element

- 1 x spare set of fuses

- User manual with trouble shooting guidance

- Installation and training at end-user site

2) Phototherapy unit, Double surface, detachable LED type

Technical Specifications:

- Heavy sturdy mobile stand phototherapy unit

- Antistatic castors, with breaks

- Double surface, detachable LED (Light Emitting Diode), surface size, approx: 0.50 x 0.75 m

- Head height adjustable, approx: 1.40 to 1.75 m

- LED’s should fully protected

- Single head, surface size, approx: 30x 60 cm

- Head height adjustable, approx: 100 to 150 cm

- Irradiance at skin level > 40uW/cm2/nm at 30 cm from light unit

- Wavelength: 400 to 500 nm, with highest intensity between 450 to 465 mm

- Integrated cumulative hour timer for unit usages & resettable therapy time

- Power requirement: 230 V/50 Hz

- Power consumption: <40 W

- Effective Light Foot Print of 500 x 250 mm to 35 mm from the light unit

- Irradiance ratio ≥ 0.4 within the effective foot print

- LED Guarantee to work for 30,000 Hours

- Device is produced by ISO 13485 Certified Manufacturer

- Device is safety certified according CE, FDA or equivalent

Supplied with:

- 1 x spare set of fuses

- User manual with trouble shooting guidance, in English

- Installation and training at end-user site

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-02 |  |  |  |
| --- | --- | --- | --- |
| Newborn Unit Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty |  |
| 1 | Resuscitator, hand-operated, neonate, 250 ml | 12 |  |
| 2 | Resuscitator, hand-operated, neonate, 500ml | 63 |  |
| 3 | Oxygen hood S and M set of 3 each, including connecting tubes | 17 |  |
| 4 | Moniter, Vital sign, NIBP, HR, SpO2, ECG, RR and Temp. | 11 |  |
| 5 | Pulse oximeter, bedside, neonatal | 8 |  |
| 6 | Syringe pump, 10,20,50 ml, single phase | 23 |  |
| 7 | ECG unit, 3 channal, portable/SET | 6 |  |
| 8 | Oxygen concentrator | 1 |  |
| 9 | Pump, suction, Foot operated | 40 |  |
| 10 | Brest Pump | 6 |  |
| 11 | Laryngoscope set, neonate | 90 |  |
| 12 | Centrifuge, hematocrit, bench top, up to 4000 RPM, including rotor | 5 |  |
| 13 | Pump, suction, portable, 220V, w/access. | 46 |  |
| 14 | Stadio Meter | 29 |  |
| 15 | ABG Analysis Machine | 1 |  |
| 16 | Portable Ultra sound | 1 |  |
| 17 | Portable X - Ray machine | 1 |  |
| 18 | Flux meter | 1 |  |
| 19 | Nebulization Machine | 39 |  |
| 20 | CPAP | 5 |  |
| 21 | Fumigation Machine | 3 |  |

1). Resuscitator, hand-operated, neonate, 250ml

Technical Specifications:

- Resuscitator for manual ventilation of neonates

- Ventilation can be done with ambient air or with oxygen

- Resuscitator can be totally disassembled, is easy to clean, disinfect and sterilize

- All parts can be autoclaved at 120°C (except O2 reserve bag)

- Manufactured from durable high-strength synthetic not requiring special maintenance or storage

- Resuscitator is supplied as a complete set in a box

- Compressible self-refilling ventilation bag, capacity, approx. 250 ml

- With non-re breathing patient valve with pressure limitation

- Intake valve with nipple for O2 tubing

- O2 reserve bag complete, capacity approx.: 800 ml

- 2 Nos of cushioned neonate size masks, translucent of each (Size “0”, approx: 50 mm)

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA 510k or equivalent.

Supplied with:

- User manual with trouble shooting guidance, in English

2). Resuscitator, hand-operated, neonate, 500ml

Technical Specifications:

- Resuscitator for manual ventilation of neonates

- Ventilation can be done with ambient air or with oxygen

- Resuscitator can be totally disassembled, is easy to clean, disinfect and sterilize

- All parts can be autoclaved at 120°C (except O2 reserve bag)

- Manufactured from durable high-strength synthetic not requiring special maintenance or storage

- Resuscitator is supplied as a complete set in a box

- Compressible self-refilling ventilation bag, capacity, approx. 500 ml

- With non-re breathing patient valve with pressure limitation

- Intake valve with nipple for O2 tubing

- O2 reserve bag complete, capacity approx.: 2000 ml

- 2 Nos of cushioned neonate size masks, translucent of each (Size “0”, approx: 50 mm)

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA 510k or equivalent

Supplied with:

- User manual with trouble shooting guidance, in English

3). Oxygen hood, Small and Medium, of each, including connecting tubes

Technical Specifications:

- Round shape or oval shape

- 3 x Small size, approx: height 22 cm, diam 25 cm

- 3 x Medium size, approx: height 18 cm, diam 20 cm

- Made of autoclavable polycarbonate

- Trauma free silicone neck, with adjustment flap

- With bilateral oxygen nozzle

- Oxygen tube of 2 m length must be provided with

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA 510k or equivalent.

Supplied with:

- 1 x spare set of tubing

- User manual with trouble shooting guidance, in English

4). Monitor, Vital Sign, NIBP, HR, Spo2, ECG, RR, and Temperature.

Technical Specifications:

- Should be reputed brand and make

- Digital Multipara monitor with 6 vital sign suitable for neonate

- It should have simultaneous monitoring of all the vital sign with some data recording

- Should have alphanumeric or feather touch keypad or touch screen

- Should have high resolution LCD screen

- For NIBP Composed of disposable or cloth cuff with inflatable bag quick-connected via tube to main unit

- Disposable or cloth is washable, strong and reinforced at both ends

- Tube length approx 60 cm

- Strip of Velcro fastening; length can be adjusted to fit around neonate upper arm

- Measuring range: up to 300 mmHg

- Minimum graduation: 1 mmHg

- Accuracy: +/- 5%

- Large LCD display readable in low light working situations, display cover durable plastic

- Displays reports: systolic, diastolic and mean pressure and heart rate

- Power consumption: 30 W

- Should have provision for external printer connection

- Should work on power line as well as battery

- Light weight design for portability

- High performance Li-ion rechargeable battery

- Battery capacity: 2 hours of normal use

- It should work on 210/270 volts 50/60Hz, Single phase

- Supplied with Powder card, 3pin molded plug

- Well Packed in Individual

- Device is produced by ISO 9001 certified manufacturer

- Manufacturer shouldhave a CEmarkedproduct

Supplied with:

- 20 x Supplied with disposable ECG electrode

- 1 x Supplied with vital sign monitoring cables and probes,

- 1 x reusable SpO2 sensors neonate, clip-on type (including connection cable)

- 1 x reusable SpO2 sensors neonate, wrap around type (including connection cable).

- 1 x set of each vital sign monitor cable and probes extra with the machine

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site

5). Pulse oxy meter, bedside, neonatal

Technical Specifications:

- Compact portable bedside pulse oxymeter with LCD display.

- Continuous monitoring of SpO2 (arterial blood oxygen saturation), pulse rate and signal strength

- Measuring range:

- SpO2: 30 to 100 %, minimal graduation 1%.

- Pulse rate: 20 to 250 bpm, minimal graduation 1 bpm.

- Accuracy SpO2: 50 to 69% (± 3%), 70 to 100% (± 2%).

- Display hows SpO2(%), HR(bpm) and signal strength bar

- Large displays readable from distance, display cover durable plastic.

- User preset of high/low alarms on SpO2 and pulse rate monitoring

- Audio visual alarm for SpO2 and pulse rate in case measurements are outside preset range

- Silencing feature for audio alarm

- Display reports system errors, probe failure and built-in battery status

- Automatic switch from mains to batteries in case of power failure

- Power requirements: 220 V/50 Hz and internal re-chargeable battery (autonomy approx 6 hrs, automatic recharge).

- Power consumption: 50 W.

- Device is produced by ISO 9001 certified manufacturer.

- Device is safety certified according CE, FDA or equivalent.

Supplied with:

- 1 x reusable SpO2 sensors neonate, clip-on type (including connection cable).

- 1 x reusable SpO2 sensors neonate, wrap around type (including connection cable).

- 1 x spare set of fuses.

- User manual with trouble shooting guidance, in English.

- Installation and training at end-user site.

6) .Syringe pump, 10, 20, 50 ml, single phase

Technical Specifications:

- Digital and self regulating volume controlled portable syringe pump

- Can be mounted on bed/wall rail or mobile pole/stand (supplied with fixation)

- Suitable for all intravenous and intra-arterial infusions

- Continuous volumetric delivery with syringes 10, 20 and 50 ml

- Open system, suitable for different brands of syringes

- Programmable, user entry: infusion volume and time or flow rate

- Rate, adjustable: 1 to 999 ml/h, steps of 1 ml/h

- Accuracy: ca 1% of total volume delivered

- With occlusion detection and alarm

- Display reports systems errors, end of infusion and built-in battery status

- Audio visual alarm with silencing feature for audio alarm

- Automatic switch from mains to batteries in case of power failure

- Power requirements: 220 V/50 Hz or internal re-chargeable battery

- Power consumption: 50 W

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA 510k or equivalent.

- Supplied with:

- User manual with trouble shooting guidance

- Installation and training at end-user site

7). ECG unit, 3 channel, portable/SET

Technical Specifications:

- 3 Channel ECG Machine with Display

- Reputed brand and make

- Should have alphanumeric or feather touch keypad

- High resolution LCD screen

- Built-in ECG Parameters measurement and interpretation

- Digital filtering to remove interference from power line, muscle tremor etc., with automatic base line correction

- Multiple print formats from integrated l thermal printer 5 / 25 / 50 mm/s (direct)

- Should have provision for external printer connection

- Should work on power line as well as battery

- Light weight design for portability

- Sensitivity: 5/ 20 / 20 mm/mV, either automatically  adjusted or manually selected

- Chart Paper :Thermo reactive – Z- folded, 210 mm wide, perforation 280 mm

- Printing Process: High-resolution thermal print head

- High performance Li-ion rechargeable battery

- Battery capacity Approx: 4 hours of normal use

- It should work on 210/270 volts 50/60Hz, Single phase

- Device is produced by ISO 9001 certified manufacturer

- Manufacturer should have a CE marked product

- Well Packed in Individual.

Supplied with:

- 2 x one bottle ECG gel

- 1 x set of bulb, clamps

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site.

8). Oxygen concentrator

Technical Specifications:

- Oxygen concentrator to provide oxygen from ambient air

- Oxygen concentration measured at the flow meter by oxygen sensing device (OSD)

- Superior grade of molecular sieve

- Maintenance free rotary proppet valve.

- Oxygen purity, approx: 90%

- Oxygen output, approx: 0 - 5 LPM

- Pressure, approx: 8 psi

- Double outlet or flow splitter for oxygen Delivery

- Oxygen tube of 1m length must be provided with

- Facility for nebulization with tube & mask

- With two humidifier bottles and two cabinet filters

- Power requirements: 220 V/50 Hz

- Power consumption: 450 W

- Device is produced by ISO 9001 certified manufacturer

- Supplied with:

- 1 x spare set of tubing

- 1 x spare set of internal and external filters (bacterial)

- 1 x spare set of fuses

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user

9). Pump Suction Foot operated

Technical Specifications:

- Foot-operated suction pump

- High performance suction pump for pharyngeal and tracheal suction

- Double acting piston pump provides a combination of large airflow and high vacuum

- See-saw movement of pedal generates suction every time one side of the pedal is depressed

- Pump chassis complete with valve diaphragms, manifold pipe, bottom cover, cylinder with draw link and valve diaphragm, piston O-ring, pedal with retaining springs, aspirating tube with angle connector and combination suction tip

- Pump can be totally disassembled, is easy to clean and disinfect

- All parts can be autoclaved at 121°C

- Vacuum, max: 400 mmHg

- Free airflow at two pumping strokes per second, approx: 30 to 40 L/min

- All parts made of high-strength, long-life materials, not requiring specific maintenance or storage

- Transparent polycarbonate collection container capacity, approx: 1 L

- Bottom cover: thermoplastic rubber

- Manifold pipe: polypropylene

- Gasket, O-rings and valve diaphragm: silicone rubber

- Piston rings: Teflon

- Foot pedal: Aluminum

- Device is produced by ISO 9001 certified manufacturer (Certificate to be submitted, further details see “Technical Provisions”)

- Device is safety certified according CE 93/42, FDA equivalent (Certificate to be submitted, further details see “Technical Provisions”)

Supplied with:

- 1x set of silicone rubber suction tubing, approx: diam. 10 mm, length 1.5 m

- 1 x angle connector and combination acetal suction tip

- 1 x spare valve diaphragms

- 1 x spare piston O-ring

- 1 x spare retaining springs,

- User manual with trouble shooting guidance, in English

- Assembly, Installation and Training at end-user site.

10). Breast Pump

Technical Specifications:

- It should be useful for extracting the excess milk, from feeding mother

- It should be made of Silicon or Rubber

- It should be contemporary design

- Device is produced by ISO 9001 certified manufacturer

- Installation and Training at end-user site

11). Laryngoscope set, neonate

Technical Specifications:

- Laryngoscope set with neonate blades

- Constituted of large hollow, cylindrical, slightly ribbed handle and a set of depressors in stainless steel

- Handle is made of either chromium-plated or stainless steel and can be opened at an extremity to insert two alkaline batteries (LR14, size C, 1.5 Volts). The other end has a stud contact which fits the various sizes and types of depressors.

- 2 x curved depressors, Macintosh type No. 1, approx 68 mm length with halogen bulb

- 2 x straight depressors, Miller type No. 0, approx 53 mm length with halogen bulb

- Presented in suitable protective plastic box

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA 510k or equivalent

Supplied with:

- 2 x spare halogen bulbs

- User manual with trouble shooting guidance

- Supplied with packed bag

12) Centrifuge, bench top, up to 4000 RPM, including rotor

Technical Specifications:

- Should be table top for quick assessment and capable of holding 8 samples at a time, and test tube capacity should be approx. 15ml

- Should be supplied with swing out hand and rotation approx. up to 4000 RPM, adjustable in increment of approx 100

- It should have time settable in min. maximum preset 99minutes

- Safety lid-lock feature and emergency lid release

- Should show the digital display of RPM and time

- Motor over heater protection and imbalance shut-off

- It should work on 200/270 volts 50/60Hz, Single phase

- Device is produced by ISO 9001 certified manufacturer.

- Device is safety certified according CE or equivalent

13). Pump, suction, portable, 220V, w/access.

Technical Specifications:

- Portable suction pump with 1or 2 litre plastic jar

- Dismantable for easy cleaning, jar autoclavable.

- Shock resistance protects from dropping

- Light, easy to carry and keep clean

- Jar with plastic cover, gasket and overflow valve

- With suction regulator, vacuum gauge

- Suction power: 15 L/min

- Vacuum, max: 600 mmHg

- Power requirement: 220 V/50 Hz

- Power consumption: 100 W

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA equivalent.

Supplied with:

- 1x set of silicone rubber suction tubing, approx: diam. 10 mm

- 1 x spare set of fuses

- User manual with trouble shooting guidance, in English

- Training and installation at end-user site.

14). Stadiometer

Technical Specifications:

- High Precision mechanism

- Contemporary stand  or fixed type design

- Should be robust construction

- Anti rust coated parts

- Should provide accurate reading

- Device is produced by ISO 9001 certified manufacturer

Supplied with:

- User manual with trouble shooting guidance, in English

- Installation and training  at end-user site

15). ABG Analysis Machine

Technical Specifications:

- Color Touch Screen - Fully Automatic Microprocessor controlled, point of Care Critical Care Analyzer with One box Cassettes, One printer Paper roll, Reusable quality controls, One Gas Bottle.

| Cassette Type | Measured Parameters | Calculated Parameters | Cost Per Test |
| --- | --- | --- | --- |
| B-Type | pH, PCO2, PO2, tHb, SO2 | HCO3, st.HCO3, BE, BEecf, BEact, BB, st.pH, cH+, tCO2, SO2, O2ct, AaDO2, P50, nCa++, Hct | 240 |
| E-Ca Type | pH, PCO2, PO2, tHb, SO2, Na+, K+, Ca++ | HCO3, st.HCO3, BE, BEecf, BEact, BB, st.pH, cH+, tCO2, SO2, O2ct, AaDO2, P50, nCa++,Hct | 260 |
| E-Cl Type | pH, PCO2, PO2, tHb, SO2, Na+, K+, Cl- | HCO3, st.HCO3, BE, BEecf, BEact, BB, st.pH, cH+, tCO2, SO2, O2ct, AaDO2, P50, nCa++,AG, Hct | 260 |
| E-GLU Type | pH, PCO2, PO2, tHb, SO2, Na+, K+, GLU | HCO3, st.HCO3, BE, BEecf, BEact, BB, st.pH, cH+, tCO2, SO2, O2ct, AaDO2, P50, nCa++, Hct | 280 |
| E-BUN Type | pH, PCO2, PO2, tHb, SO2, Na+, K+, BUN | HCO3, st.HCO3, BE, BEecf, BEact, BB, st.pH, cH+, tCO2, SO2, O2ct, AaDO2, P50, nCa++, Hct | 280 |
| B-LAC Type | pH, PCO2,PO2, tHb, SO2, LAC | HCO3, st.HCO3, BE, BEecf, BEact, BB, st.pH, cH+, tCO2, SO2, O2ct, AaDO2, P50, nCa++, Hct | 280 |

- Measured Parameters: Machine measures parameters directly with Optical Fluorescence and Reflectance Technology

- Calculated Parameters:  Parameters come from “Measured Parameters” with derivative formulas

- tHb - Total Hemoglobin                         SO2- Oxygen Saturation

- GLU- Glucose                                       BUN- Blood Urea Nitrogen

Lac- Lactate

- HCO3- Actual Bicarbonate                   st.HCO3- Standard Bicarbonate

- BE- Base Excess                                  BEecf - Base Excess Extra Cellular Fluid

- BEact - Base Excess Actual                 BB- Buffer Base

- St.pH - Standard pH                             cH+- Hydrogen Ion Concentration

- tCO2 - Total CO2                                 SO2- Oxygen Saturation

- O2ct - Oxygen Content                        AaDO2- Alveolar-arterial oxygen concentration

- P50                                                       nCa++- St. Ionized Calcium (pH=7.4)

- AG- Anion Gap                                     Hct - Hematocrit

Features:

- Storage of Consumables  at normal room temperature except BUN (2-8°C)

- Automatic sample aspiration

- Portable (light weight and Mains cum  battery operated with 8 Hrs back-up)

- Reusable Electronic Quality control

- Barcode scanner for quality control information and can be conveniently entered using the on-board bar code scanner

- On Board Thermal printer

- Display of results on in-built  active color touch screen

- 200 patient data

- RS 232 interface Infrared. ASCII

- CE, UL, IVDD Compliant

- FDA approved

- Barometric Pressure : 300–800 mmHg

Operating Parameters:

- Minimum sample size: 125 ul

- Sample type: whole blood, serum, plasma

- Sample application: Syringe, capillary or Micro sampler

- Sample Input: Automatic Aspiration

- Analysis time (time to result) < 120 sec.

- Ambient temp: 10°C – 32 °C

- Relative Humidity: 5% - 95% (non-condensing)

Input Parameters:

- Patient temp, Hemoglobin type ,p50,  FlO2,RQ

- Operation ID. Patient ID No., Patient Sex

- Mean Corpuscular Hemoglobin Concentration (MCHC), Sample type:  art/ven/mix, ven/cap/cord/CPB

- User - defined field,3 User- defined fields:

- Puncture site: left radial, right radial, left brachial, right brachial, left femoral, right femoral, cord, scalp

Data Management:

- Print out: Built-in Thermal printer

- Interface: RS 232 C, ASCII or ASTM format

- Quality Control Data & standard reference cassette report: 1 month of 3 levels stored onboard.

Security:

- Password: up to 2 levels – User and administrator

- QC lock out: User definable, Number of levels / selectable time period, Operator ID.

Electrical & Gas supply:

- Voltage / Frequency: 100 – 250 VAC / 16 VDC, 2.8 A / 50 – 60 Hz.

- Power Consumption (Maximum): 110 VA

- Battery operation: Up to 8 hrs. (Rechargeable battery pack)

- Calibration Gas: Disposable small cylinder

Dimensions / Weight:

- Height x Width x Depth: 124 mm x 362 mm x 230 mm / 4.7 in x 14.2 in x 9.1 in

- Weight: 4 Kg (9 lbs) w/o battery 5 kg (11 lbs) with battery.

Supplied with:

- Well Packed in Individual

- User manual with trouble shooting guidance

- Installation and Training at end-user site

16).Portable Ultra Sound with ECHO (USG)

Technical Specifications:

Electric Power

- Voltage 100 – 240 V AC

- Frequency 50/60 Hz

- Power: Max 130 VA with Peripherals Console Design

- Portable type

- Integrated HDD (20 GB) Approx.

User Interface

Display Screen

- Approx. 8’’ inch high-resolution color LCD display

- Interactive Dynamic software menu

- Open Angle Adjustable approx. 0 to 120°

- Brightness Adjustment

System Overview

Applications

- Cardiology

- Neonatal and Pediatric

Scanning Methods

- Electronic Convex

- Electronic Linear

- Cardiac probe

Operating Modes

- B, Mode

- M, Mode

- B+M, Mode

- 2B,

- 4B,

- ABD, OB, GYN, CARD, URO, Etc.

- B Mode

- Frequency

- Grey Map

- Dynamic Range

- Focus Position

- Edge Enhance

- Focus Number

- M Mode

- Grey Map

- Dynamic Range

- Sweep Speed

- Display Format

- Edge Enhance

- Gain: 0-10 dB, 1 dB step

Supplied with:

- 2 x one bottle USG gel

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site

- B/W Thermal printer.

17). Portable X - Ray machine

Technical Specifications:

- Carry them very conveniently for use anywhere in hospital and also in :

- Pediatric, Neonatal ward and Trauma and emergency medicine

- Sports medicine.

- Any remote location.

- Certified by the AERB (Atomic Energy Regulatory Board) for radiation safety and conform to relevant BIS norms. This assures you of radiation safety during operation and consistency in performance.

- Solid state 23 step electronic timer from 0.06 to 6 sec.

- Light beam collimator.

- Tape for measuring film focus distance.

- X-Ray Generator

Rating			50mA at 50 KV from 0.06 to 6.0 secs.

40mA at 65 KV from 0.06 to 6.0 secs.

35 mA at 75 KV from 0.06 to 6.0 secs.

30mA at 85 KV from 0.06 to 6.0 secs.

- Exposure time		Solid state 23 step electronic timers from 0.06 to 6 sec.

- X-ray tube		Stationary anode, I.B mm X 1.8 mm focus.

- Protection		Built in over voltage protection and indication.

- Collimator		Light beam collimator for centering and radiation protection.

- L V compensation	Line voltage compensation from 210 to 250 volts.

- Power			Single Phase 230 volts, 15 Amps, 50 Hz. Requirements

Supplied with:

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site

18). Flux Meter for Neonates:

Technical Specifications:

- Reputed brand and make

- It should work on 210/270 volts 50/60Hz, Single phase

- Device is produced by ISO 9001 certified manufacturer.

Supplied with:

- Well Packed in Individual

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site.

19). Nebulization Machine for Children

Technical Specifications:

- It should be High Precision mechanism

- It should  provide nebulization  from ambient air

- Contemporary design of round or Square

- It should be maintenance free

- Single outlet  for nebulization delivery

- Nebulization  tube of 1m length must be provided with tube &mask

- Well Packed in Individual

- Supplied with Powder card,  molded plug

- Power requirements: 220 V/50 Hz

- Device is produced by ISO 9001 certified manufacturer

- Supplied with:

- 1 x spare set of tubing & mask

- 1 x spare set of fuses

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site.

20). CPAP for Neonate & Children

Technical Specifications:

- It should be High Precision mechanism

- It should be contemporary design

- It should be maintenance free

- Supplied with suitable compressor, standard accessories for smooth functioning the CPAP.

- Must be provided with tube &mask

- Supplied with Powder card, molded plug

- Power requirements: 220 V/50 Hz

- Device is produced by ISO 9001 certified manufacturer

- Supplied with:

- 1 x spare set of tubing & mask

- Well Packed in Individual

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site

21). Fumigation machine for children wards and SCNU

Technical Specifications:

- It should be contemporary design

- It should be maintenance free

- Should  have single outlet delivery system

- must be provided with tube &mask

- Power requirements: 220 V/50 Hz

- Device is produced by ISO 9001 certified manufacturer

- Supplied with:

- Well Packed in Individual

- Supplied with Power card,  molded plug

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site.

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-03 |  |  |
| --- | --- | --- |
| Blood Bank Equipments |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty |
| 1 | Blood Bank Refrigerator (2 o C - 8 o C) | 9 |
| 2 | Water Bath | 13 |
| 3 | Precision Incubator | 13 |
| 4 | Binocular Microscope | 11 |
| 5 | Table Top Centrifuge | 10 |

- Blood Bank Refrigerator vertical type (2°C - 8°C)

Technical Specifications:

- Blood bank refrigerator Capacity is Approx 60 bags.

- Exterior Galvanized MS with white powder coated body and dark blue panel cover with Interior cabinet made of AISI SS 304.

- Temperature: 4°C +/-1°C.

- Insulation CFC - Free light density poly therein foam insulation

- Closure: Durable magnetic rubber gaskets for leak proof closure

- Main door: Left hinged double pane glass door

- Heater Line: Hot air over the front door to prevent moisture condensation

- Refrigeration: hermetically sealed compressor with non CFC. HCPC eco friendly refrigerant (R- 134A)

- Compressor: Heavy duty hermetically sealed compressor

- Cooling positive force air circulation for uniform cooling throughout the cabinet

- Condenser By means of air cooled compressor aided by continuous rated motor fans.

- Control system: Microprocessor based compressor digital temperature controller with digital display of set and process temperature.

- UPS backup built in battery up for alarm system and temperature display during

- power failure

- Probe: Fast response PT-100 RTD sensor for accurate temperature measurement

- Safety: Slow blow Fuse & OLP’s

- Smooth moving caster wheels

- Over load protection compressor with internal /external over load protector.

- Alarm audiovisual alarm for temperature deviations & power failure.

- Temperature recorder built in microprocessor based weekly circular chart temperature record with chart paper for one year.

- Stabilizer Automatic voltage stabilizer with time delay chart

- Calibration: Temperature calibration certificate traceable to National Standards (ETDC) or National Physical Laboratories (NPL).

- Temperature range 5°C above ambient to 65°C.

- Temperature control +/-0.5°C

- Display & Setting +/-0.5°C

- Temperature controller fully microprocessor, controlled with soft touch key.

- Temperature display should be digital

- Temperature sensor should have 100RTD (PT-100) with suitably sheathed

- Relay electronic, heater capacity 1 or 2 kv.

- Construction: Double walled outer MS powder coated inner & lid stainless steel SS304 non magnetic.

- Dimension: 300x250x 175mm

- Power: 230+/- 10%, 50Hz Single phase.

- BBR – UPS, Connection should be provided directly when ups is failed, but BBR should work

2. Water bath

Technical Specifications:

- Water bath should be is a tabletop should have minimum bath capacity of 10 to 20 litres,

- Exterior Galvanized MS with white powder coated body and Interior cabinet made of AISI stainless steel inner structure SS 304

- There should be temperature control of 4ºC C to 65ºC or better, with digital display of temperature, reliably thaw

- Unit must be fitted with microprocessor based digital temperature Indicator cum controller.

- Circulation of water inside the closed chamber to get the uniform temperature, high quality rust and corrosion resistant stainless steel reservoir

- Automatic and manual reset after power failure.

- The unit must be provided with the compatible non-corrosive lid for internal chamber and the equipment must be leakage proof.

- There should be provision of replacement of water.

- Should be ensuring safely of with ±0.2ºC stability.

- It should work on 115/240 V~ -50/60 Hz, in single phase

- Device is produced by ISO 9001 certified manufacturer

- Should have certificate of National Physical Laboratories (NPL)

3. Precision Incubator

Technical Specifications:

- Incubator should be stainless steel inner structure.

- Exterior Galvanized MS with white powder coated body and Interior cabinet made of AISI stainless steel inner structure SS 304

- There should be temperature control of 4ºC C to 65ºC or better, with digital display of temperature

- Unit must be fitted with microprocessor based digital temperature Indicator cum controller

- Digital control digital set and readout.

- Should have forced convection uniform temperature control

- Electronic control with safety backup

- Should have optional humidifier

- Should maintain uniforms temperature in the chamber.

- Inner tempered glass door with silicon packing

- Should have easy sample view without door opening stainless steel chamber.

4. Binocular microscope

Technical Specifications:

- Body: Should be inter changeable identified binocular body 360o   rotatable head

- Eye pieces should be high quality 10x wide angle anti fungus field eye piece

- Objectives: Par focal anti fungus coated 4x 10x 40 and 100x (full immersion) with plan automatic correction

- Optical system should be achromatic objectives

- Stage: Horizontal mechanical stages preferable 140,440mm with fine vernier graduations

- Design with convenient coaxial adjustment for slide manipulation preferably

- Through 30x70mm.

- Substage: Abbe condenser focusable continuously variable iris diaphragm

- Illuminator with illumination

- Finish: A durable textured acid resistant finished.

- It should be safety and convenience over temperature and current protection.

- Should have safety and adjustable sliding rack,

- It should work on 200/240 V~ -50/60 Hz,

- Device is produced by ISO 9001 certified manufacturer.

- Should have certificate of National Physical Laboratories (NPL)

5. Table Top Centrifuge

Technical Specifications:

- Should be table top for quick assessment and capable of holding 8 samples at a time, and test tube capacity should be approx. 15ml

- Should be supplied with swing out hand and rotation approx. up to 4200 RPM, adjustable in increment of 100

- It should have time settable in min. maximum preset 99minutes

- Safety lid-lock feature and emergency lid release

- Should show the digital display of RPM and time

- Motor over heater protection and imbalance shut-off

- It should work on 200/270 volts 50/60Hz, Single phase

- Device is produced by ISO 9001 certified manufacturer.

- Device is safety certified according CE or equivalent.

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-04 |  |  |  |
| --- | --- | --- | --- |
| Diesel Generator Sets - 5KVA & 25 KVA |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs: |
| 1 | Generator set 25 KVA | 5 |  |
| 2 | Generator 5 KVA | 25 |  |

| 1. Generator (DG) Set 25KVA |  |
| --- | --- |
| Technical Specifications: - 25KVA DG SET comprising of ENGINE, with standard accessories coupled with Alternator with Diesel Tank, Starting battery with leads, AMF Control panel - Should be reputed brand and make and should be Silent D.G Set with Acoustic Enclosure - 25KVA DG SET comprising of ENGINE, with standard accessories coupled with Alternator withDiesel Tank, Starting battery with leads - Supplied with standard AMF panel of fully automatic control, and Comprising of Alternator Stamford/KEC/CG - It should CPCB, Approved Acoustic Enclosure - Rating should be in KVA, PF Brushless - Digital Type Voltmeter, Ammeter, Hertz Meter, Kwh - MCB with Acoustic Enclosure as Per CPCB Norms - Water-cooled or Air cooled, 36.9 BHP Four Cylinders or two cylinder - SafetyKeylocktoprevent unauthorizeduse - Device is produced by ISO 9001 certified manufacturer Supplied with: - supplied with all standard accessories and fixing and fitting with laying, - User manual with trouble shooting guidance, in English Installation and Training at end-user site |  |

TECHNICAL SPECIFICATIONS

Scope of Work

Scope of the work shall include Supply, Erection, Testing and Commissioning of 25 KVA, SILENT Diesel Generator Set with Central Pollution Control Board approved Acoustic Enclosure comprising of Diesel Engine with standard accessories coupled to Alternator mounted on Channel Iron Base Frame with Vibration Free Pads, complete with Fuel Tank having sufficient capacity for minimum 8 Hours of operation with Auto Main Failure (AMF) Control Panel and Batteries with the following features given below.

Make

Engine	: Kirloskar Electric / Asok Leyland / Cummins / Mahindra / Equivalent

Alternator	: KIRLOSKAR ELECTRIC / CROMPTON GRAVES / STANFORD / EQUIVALENT

Engine

The Engine shall be of Radiator Cooled, Four Stroke, Multi-Cylinder confirming to relevant Standards.

Battery Charging Alternator shall be coupled with the Engine.

Lube Oil Pressure Gauge, Water Pressure Gauge, Water Temperature Gauge and Oil Temperature Gauge shall be provided in the Engine.

Additional 230V Battery Charger shall be provided external to the engine.

Engine shall be filled with first fill of (fresh) Lube Oil and Coolant.

Alternator

Alternator shall be of Brushless, Self-Excited type rated at 25 KVA.

Terminal Voltage of the Alternator shall be 3 Phase, 4151% Volts, 50 Hz, 0.8 Power Factor, Four Wire.

Insulation of the Alternator shall be Class-H.

Voltage of the Alternator shall be self-regulated through AVR.

Mounting

Engine and Alternator shall be mounted on a common MS fabricated base frame with AVM Pads.

AMF Control Panel

The Control Panel shall be fabricated with minimum 16 gauge CRCA sheet and painted with powder coated paint. Suitable size of Aluminum Buses shall use for incoming and outgoing termination.

The DG set shall work on Auto Main Failure (AMF) with all accessories such as Contactors having sufficient rating, AMF Relays etc. complete. AMF Relays etc. complete. AMF control shall be inbuilt, Factory integrated along with the Generator.

The operation, monitoring and control of the system shall be controlled with Micro-processor based Controller installed in the Control Panel.

CT (Current Transformer) provided for measuring current shall be of Class 0.5 and shall be of cast resin type. The CT’s provided shall be submitted with the test certificate from State Electrical Inspectorate.

Indication Lamps shall be provided in the control panel for ‘Load On’ and ‘Running’.

Suitable rating of MCCB with over load and short circuit protection shall be mounted in the Control Panel for the isolation of loads.

MCB shall be provided for the isolation / protection of control circuit (No fuses shall be used).

The AMF Control panel shall have the following functional facilities

- Emergency Stop

- Fault Reset

- Auto Start

- Manual Start

- Meters displaying the parameters

- Voltage

- Current

- Frequency

- Speed (RPM)

- Oil Pressure

- Water Pressure

- KVA

- KW

- LED Indication for the faults

- Fail To Start

- Over Load

- Over Speed & Under Speed

- Low Oil Pressure

- Engine temperature High

- Battery Charging Alternator Fail

- Engine Hour Counter

Vendors

Vendors List of the switchgear and other accessories are given below.

MCB & MCCB		: SIEMENS

Contactor			: SIEMENS

CT			: KAPPA / INTRANS / Equivalent

Indication Lamps		: TECKNIK / SIEMENS / L&T

Meters			: COZERVE / RISHAB /AE

Acoustic Enclosure

Acoustic Enclosure shall be made of CRCA sheet having minimum 2 mm thickness and the walls of the enclosure insulated with fire retardant foam. Sound level at 1 Mtr shall be ≤75 dB. The DG set shall meet all the specifications/norms stipulated by Ministry of Environment and Forests, Central Pollution Control Board (CPCB) and State Electrical Inspectorate etc. Necessary Air Louvers shall be provided in the enclosure.

Battery

Battery shall be Sealed Maintenance Free type. Make and Technical Data of the Battery shall be mentioned in the offer.

Fuel Tank

Fuel Tank shall have sufficient capacity for minimum 8 Hours of operation of the DG Set. Fuel Gauge shall also be provided for the indication of the level of the fuel. Breather to be provided on the top of the tank.

Installation, Testing & Commissioning

Supply, Installation, Testing & Commissioning of the whole system shall be in the scope of the Supplier.

Minimum 30Ltr of Diesel shall be filled by the supplier for the commissioning the DG at free of cost.

Approval

Approval from statutory bodies shall be arranged by HLL

Warranty

1) The Engine, Alternator, Battery and other accessories shall be warrantee for a period of 12 Months from the date of commissioning. Necessary Certificates of Warranty shall be submitted along with the DG Set.

2) Supplier has to provide periodic Preventive Maintenance service.

3) Breakdown Maintenance support, for entire warranty period, shall be provided within 8 Hours ofintimation of breakdown.

4) Preventive Maintenance service& Breakdown maintenance service support from the part of DG supplier includes:

- Periodic change of Engine oil, filter etc at the cost of supplier including consumables like Engine oil, filter etc.

- Replacement of all requisite spares and accessories.

- All other consumables as required for the service.

- Battery consumables like distilled water etc.,  shall be under the scope and cost of DG supplier.

- Transportation, loading unloading of consumables, spares and accessories brought for maintenance service and service charges..

Annual Comprehensive Maintenance Contract (ACMC)

Tenderer shall extend Annual Comprehensive Maintenance Contract for the whole system for a period of FIVE years after the completion of Performance Warranty period. Supplier has to provide service support within 8 Hours. Rate of Comprehensive Annual Maintenance Contract shall be quoted in the Format attached as ANNEXURE.

Special Conditions:

- Make and Model Number of Diesel Engine and Alternator shall be specifically mentioned in the Tender. Catalogue of the same shall also be attached with the offer.

- BHP of the Diesel Engine shall be mentioned in the Tender and the same shall be taken for the calculation of power conversion.

- Rate may be quoted in the specific Price Bid Format attached

- Pre-Shipment Inspection and Testing to be conducted by the supplier in the presence of Engineer in charge of HLL.

- Test Certificates of Engine and Alternator shall be supplied along with the DG Set.

- The DG set shall be placed on the foundation by the supplier.

- All the Unloading and Leading works of the DG Set at the site shall be in the scope of the supplier.

- A competent Technical Person from the supplier shall be present at the time of Trial Run and for the successful Commissioning of the DG Set.

- Approval of Equallant make shall be strictly under the preview of HLL and HLL can accept / reject any makes coming under equallent without assigning any reason.

- Tenderer shall provide and co-ordinate Service/Technical support for Engine, Alternator and other accessories during the period of Performance Guarantee and AMC Period.

| Generator (DG) Set 5 KVA | 5 Nos. |
| --- | --- |
| Specification |  |
| Technical Specifications: - Should be reputed brand and make - 5KVA DG SET comprising of ENGINE, with standard accessories coupled with Alternator with - Diesel Tank, CPCB Approved Acoustic Enclosure - Type Air cooled - Rated Power HP 9 - Speed RPM 3000 - Fuel Diesel - No of phase 1 - Rating KVA 5 - Voltage 230 V - Frequency 50 - Power factor 1 - Insulation class H Supplied with: - Well Packed in Individual - The equipment should have lockable castors - Device is produced by ISO 9001 certified manufacturer - User manual with trouble shooting guidance, in English - Installation and training at end-user site. |  |

Scope of Work

Scope of the work shall include Supply, Erection, Testing and Commissioning of  5 KVA, Single Phase, SILENT Diesel Generator Set with Central Pollution Control Board approved Acoustic Enclosure comprising of Diesel Engine with standard accessories coupled to Alternator mounted on Channel Iron Base Frame with Vibration Free Pads, complete with Fuel Tank having sufficient capacity for minimum 8 Hours of operation with Auto Main Failure (AMF) Control Panel and Batteries with the following features given below.

Make

Engine		: HONDA / Kirloskar / Mahindra / Equivalent

Alternator	: KIRLOSKAR ELECTRIC COMPANY / Honda / SINCRO / NSM / Equivalent

Engine

The Engine shall be of Air Cooled type to relevant Standards.

Battery Charging Alternator shall be coupled with the Engine.

Lube Oil Pressure Gauge, Water Pressure Gauge, Water Temperature Gauge and Oil Temperature Gauge shall be provided in the Engine.

Additional 230V Battery Charger shall be provided external to the engine.

Engine shall be filled with first fill of (fresh) Lube Oil and Coolant.

Alternator

Alternator shall be of Brushless, Self-Excited type rated at 5 KVA.

Terminal Voltage of the Alternator shall be Single Phase, 230 Volts, 50 Hz, 0.8 Power Factor, Insulation of the Alternator shall be Class-F/B.

Voltage of the Alternator shall be self-regulated through AVR.

Mounting

Engine and Alternator shall be mounted on a common MS fabricated base frame with AVM Pads.

AMF Control Panel

The Control Panel shall be fabricated with minimum 16 gauge CRCA sheet and painted with powder coated paint. Suitable size of Aluminum Buses shall use for incoming and outgoing termination.

The DG set shall work on Auto Main Failure (AMF) with all accessories such as Contactors having sufficient rating, AMF Relays etc. complete.

The operation, monitoring and control of the system shall be controlled with Micro-processor based Controller installed in the Control Panel.

CT (Current Transformer) provided for measuring current shall be of Class 0.5 and shall be of cast resin type. The CT’s provided shall be submitted with the test certificate from State Electrical Inspectorate.

Indication Lamps shall be provided in the control panel for ‘Load On’ and ‘Running’.

Suitable rating of MCCB with over load and short circuit protection shall be mounted in the Control Panel for the isolation of loads.

MCB shall be provided for the isolation / protection of control circuit (No fuses shall be used).

The AMF Control panel shall have the following functional facilities

- Emergency Stop

- Fault Reset

- Auto Start

- Manual Start

- Meters displaying the parameters

- Voltage

- Current

- Frequency

- Speed (RPM)

- Oil Pressure

- KW

- LED Indication for the faults

- Fail To Start

- Over Load

- Over Speed & Under Speed

- Low Oil Pressure

- Engine temperature High

- Battery Charging Alternator Fail

- Engine Hour Counter

Vendors

Vendors List of the switchgear and other accessories are given below.

MCB & MCCB		: SIEMENS / Equivalent

Contactor			: SIEMENS / Equivalent

CT			: KAPPA / INTRANS / Equivalent

Indication Lamps		: TECKNIK / SIEMENS / L&T / Equivalent

Meters			: COZERVE / RISHAB / Equivalent

Acoustic Enclosure

Acoustic Enclosure shall be made of CRCA sheet having minimum 2 mm thickness and the walls of the enclosure insulated with fire retardant foam. Sound level at 1 Mtr shall be ≤75 dB. The DG set shall meet all the specifications/norms stipulated by Ministry of Environment and Forests, Central Pollution Control Board (CPCB) and State Electrical Inspectorate etc. Necessary Air Louvers shall be provided in the enclosure.

Battery

Battery shall be Sealed Maintenance Free type. Make and Technical Data of the Battery shall be mentioned in the offer.

Fuel Tank

Fuel Tank shall have sufficient capacity for minimum 8 Hours of operation of the DG Set. Fuel Gauge shall also be provided for the indication of the level of the fuel. Breather to be provided on the top of the tank.

Installation, Testing & Commissioning

Supply, Installation, Testing & Commissioning of the whole system shall be in the scope of the Supplier.

Minimum 15 Ltr of Diesel shall be filled by the supplier for the commissioning the DG at free of cost.

Approval

Approval from statutory bodies shall be arranged by HLL

Warranty

The Engine, Alternator, Battery and other accessories shall be warrantied for a period of 12 Months from the date of commissioning. Necessary Certificates of Warranty shall be submitted along with the DG Set. Supplier has to provide service support within 8 Hours from the date of commissioning.

Annual Comprehensive Maintenance Contract (ACMC)

Tenderer shall extend Annual Comprehensive Maintenance Contract for the whole system for a period of FIVE years after the completion of Performance Warranty period. Supplier has to provide service support within 8 Hours. Rate of Comprehensive Annual Maintenance Contract shall be quoted in the Format attached as ANNEXURE.

Payment Terms

Supply Part

- 70% within 30days on documentary proof of invoiced items on Pro-Rata basis, Goods Receipt Certificate designated Hospitals in Himachal Pradesh State and Warranty Certificates.

- 20% after Erection, Testing, Commissioning and Handing Over of the items.

- 10% after Erection, Testing, Commissioning and Handing Over of the complete system against Performance Bank Guarantee valid for a period of ONE year.

Service Part

- 100% after Erection, Testing, Commissioning and Handing Over of the complete system at designated Hospitals in Himachal Pradesh State against Performance Bank Guarantee valid for a period of ONE year.

Special Conditions:

- Make and Model Number of Diesel Engine and Alternator shall be specifically mentioned in the Tender. Catalogue of the same shall also be attached with the offer.

- BHP of the Diesel Engine shall be mentioned in the Tender and the same shall be taken for the calculation of power conversion.

- Rate may be quoted in the specific Price Bid Format attached

- Pre-Shipment Inspection and Testing to be conducted by the supplier in the presence of Engineer in charge of HLL.

- Test Certificates of Engine and Alternator shall be supplied along with the DG Set.

- The DG set shall be placed on the foundation by the supplier.

- All the Unloading and Leading works of the DG Set at the site shall be in the scope of the supplier.

- A competent Technical Person from the supplier shall be present at the time of Trial Run and for the successful Commissioning of the DG Set.

- Approval of Equivalent make shall be strictly under the preview of HLL and HLL can accept / reject any makes coming under equivalent without assigning any reason.

- Tenderer shall provide and co-ordinate Service/Technical support for Engine, Alternator and other accessories during the period of Performance Guarantee and CMC Period.

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-05 |  |  |
| --- | --- | --- |
| Hospital Equipments |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty |
| 1 | Infant meter, plexus, 3½ft/105cm | 4 |
| 2 | Light, examination, mobile,220-12V | 2 |
| 3 | Spot Lamps | 9 |
| 4 | Crash Cart | 9 |
| 5 | Medicine Trolly | 10 |
| 6 | Basin, kidney, stainless steel, 825ml | 24 |
| 7 | Tray,dressing,ss,300x200x30mm | 23 |
| 8 | Glucometer with Dextrostix | 18 |
| 9 | X-Ray Viewing Box | 60 |
| 10 | Hub Cutter, syringe | 40 |
| 11 | E.Weighing Scale Digital - 100 Kg. | 29 |
| 12 | Scale, Baby, Electronic, 10 kg <5kg> | 67 |
| 13 | Weighing Scale Digital - 10Kg (Resolution up to 2 grams) | 2 |

1). Infant meter, Plexi, 100cm

Technical Specifications:

- Portable baby/infant length-height measuring system

- Measures laying length of neonates and babies

- No need for calibration as all parts have prefixed position

- Reads in centimeters and inches

- Minimum graduation: 1 mm

- Long-lasting hard-wearing ruler/graduation is fully integrated with device

- Measuring slide/wedge glides smoothly and close via ruler, avoiding reading parallax

- Measuring slide/wedge wobbles max 2 mm, over full length

- No sharp edges or corners

- Length, measurement range, approx: 100 cm

- Head/footplate, board and slide/wedge made of quality laminated wood or plastic

- Device is produced by ISO 9001 certified manufacturer.

Supplied with:

- Well Packed in Individual

- User manual with trouble shooting guidance, in English

- Installation and training at end-user site.

2). Light, examination, mobile, 220V-12V

Technical Specifications:

- Mobile light for medical examination

- Stand with anti-static swivel castors

- Articulated arm, spring loaded, with on/off switch and integrated transformer

- Halogen bulb approx: 12V/20W

- Light intensity approx: 20000 Lux

- Natural white light: colour temperature 4000 K

- Reflector adjustable for positioning

- Power requirements: 220 V/50 Hz

- Power consumption: 50 W

- Device is produced by ISO 9001 certified manufacturer.

Supplied with:

- 1 x spare halogen bulb

- User manual with trouble shooting guidance, in English

- Assembly and installation at end-user site.

3). Spot Lamp for Emergency

Technical Specifications:

- It should be Flexible gooseneck arm

- Integrated Transformer

- Stand with anti-static swivel castors.

Supplied with:

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site.

4). Crash Cart

Technical Specifications:

- Overall size shall be more than 900mm L x 500mm W x 1500 mm H.

- The crash cart should be made of SS grade tabular frame and SS sheet grade.

- Should have push handles

- Should have SS Shelves, Six color removable bins & two polystyrene lockable storage units with  drawers each.

- Facility to carry any Monitors or Defibrillators, etc. On open areas at top and bottom shelves

- Should have Stainless rods fixed with.

- Two accessory mounting brackets to mount accessories anywhere without the need of pre-threaded holes.

- Crash cart should be mounted on 12.5cms dia non-rusting swiveling castor wheels, two having locking arrangement.

- It should be robustness in quality.

5). Medicine Trolley

Technical Specifications:

- Overall size shall be more than 35cm W x 45cm L x 5cm Depth, Height 750cm.

- It should be made of best quality SS sheet

- Should be light weight

- It should be robustness in quality

- It should have the provision for hold on hands.

6). Basin, kidney, stainless steel, 825ml

Technical Specifications:

- Stainless steel kidney basin

- Capacity, approx: 825 ml

- Austenitic stainless steel, composition: 8 to 10% nickel, 18 to 20% chromium

- Thickness steel: 0.8 mm

- Length, approx: 250 mm

- Width, approx: 140 mm

- Height, approx: 40 mm

- Device is produced by ISO 9001 certified manufacturer.

7). Tray, dressing, SS, 300x200x30mm

Technical Specifications:

- Stainless steel dressing tray

- Rectangular with rounded corners

- Austenitic stainless steel, composition: 8 to 10% nickel, 18 to 20% chromium

- Thickness steel: 0.8 mm

- Length, approx: 300 mm

- Width, approx: 200 mm

- Height, approx: 30 mm

- Device is produced by ISO 9001 certified manufacturer.

8). Glucometer with Dextrostic.

Technical Specifications:

- Should be reputed brand and make

- Glucometer with 100 Strips.

- Long lasting Battery operated.

- User manual with trouble shooting guidance, in English

9). X-Ray Viewing Box

Technical Specifications:

- It will be made of thin metal body frame,  integrated ON/OFF switch,

- Provision for view one x – ray film on it

- It will be smooth finished body with approx. 2 meter power cable

- It should work on 200/270 volts 50/60Hz, Single phase

- Device is produced by ISO 9001 certified manufacturer.

Supplied with:

- Well Packed in Individual

- Supplied with Powder card, 3pin molded plug

- User manual with trouble shooting guidance

10). Hub cutter, syringe.

Technical Specifications:

1. Material used will have

- The cutting blade of the cutter will be made of stainless steel.

- The thickness of the blade will not be less than 0.5 mm with one side sharpen to cut minimum of   5,000 cuts.

- All other metal if used in the hub cutter will be made of galvanized carbon steel which will withstand rusting even in rough weather with high humidity coastal region for a minimum of two year (IS 277 of 2003 – galvanized steel sheet specification).

- The spring if used in the hub cutter will be made of steel wire for mechanical spring (oil harden and tampered steel wire)

- The sharp container should be made of Polypropylene (PP) of not less than 4 mm thick and should be white/ translucent. The material should be autoclavable plastics, puncture resistant, high drop impact strength, material should be non-toxic and pyrogen free, the material will not    wear out with normal usage.

2. Shape:

- The shape of the container will be such that the lid is tightened with the body with     a twist and cannot be opened by pulling the lid and the body in the opposite direction. The     container containing the cut hub along with needle of various size will be ¾ full when filled with     200 cut hub with needles.

3. Capacity range:

- Able to destroy from hub of AD syringe with needles from 18-28 gauge diameter and from 10-25 mm length and fitted with all types of needle fittings, eg, fixed luer lock, luer ship of snap-on

- Resistance to piecing:

- Container shall accept 200 cut needles along with the hub when it is ¾th     full (filling line). Needles shall not penetrate container walls when filled up to the filling line. The     average of forces needed to penetrate samples shall be not less than 15 N.

- Tamper-proofing:

- When contaminated needles are being dropped into the containers, it shall remain sufficiently sealed so as to prevent a hand from entering. Needle should not protrude from     the container when the box contains 200 needles with hub.

- Capacity of the container:

- Have sufficient capacity to hold 200, 18-28-gauge, 10-25 mm long needles with hub.

- Color:

- The container shall be translucent white in color and will allow visual inspection from outer surface to determine when the container is full

- Technique:

- The device shall be designed to facilitate the location and entry of the needles so that the AD syringe is cut from hub by having:

- A non-critical angle of the needle relative to the entry point.

- an entry target area which is recessed, concave and diameter is such that it allows all type of AD syringe available in the market to cut from hub

- Handling:

- Able to be carried in one hand and portable.

- Durability:

- The performance and safety of the device shall not be compromised by Exposure to ambient conditions of +43C, 90% relative humidity for one week

Supplied with:

- Well Packed in Individual

- Device is produced by ISO 9001 certified manufacturer

- User manual with trouble shooting guidance, in English

- Installation and training at end-user site

11). Electronic Weighing Scale digital 100 kg

Technical Specifications:

- Electronic weighing scale for children up 18 yrs

- Weighing capacity should be approx. 0.5kg  to 100 kg

- With tare function

- On switch and auto-off

- Display in kg

- Reading time max 5 seconds

- High Precision mechanism

- Should have fine zero adjustment

- Contemporary round or Square in design

- Should be robust construction & portable.

- Anti rust coated parts

- Wide window for clear & better visibility

- Should provide accurate reading

- Device is produced by ISO 9001 certified manufacturer.

Supplied with:

- User manual with trouble shooting guidance, in English

- Installation and training at end-user site.

12). Scale, baby, electronic, 10 kg <5g>

Technical Specifications:

- Electronic scale for weighing new born babies

- Measuring range 0 to approx 10 kg

- Minimum graduation: 10 g

- With tare function

- On switch and auto-off

- Auto-calibration with each switch-on

- Large LED display readable in low light working situations, display cover durable plastic

- Display in kg and lbs, easy switch between kg and lbs

- Reading time max 5 seconds

- Zero weighing adjustment

- Freeze reading feature

- Smooth surface/finishing allows for easy cleaning/disinfection.

- All vital parts made of rust proof materials

- Horizontal leveling with height adjustable feet

- Splash proof and shock resistant light-weight body

- Power requirements: 220 V/50 Hz

- Device is produced by ISO 9001 certified manufacturer

- Device is safety certified according CE 93/42, FDA 510k or equivalent

Supplied with:

- 1 x spare set of fuses

- User manual with trouble shooting guidance, in English

- Training and installation at end-user site

13). Electronic weighing machine for Neonates:

Technical Specifications:

- Weighing scale for neonate electronic, this can weigh up to 10kg,

- with resolution up to 2grams, which can be used to weigh diapers of the newborn also.

- Supplied with:

- Well Packed in Individual

- User manual with trouble shooting guidance, in English

- Installation and Training at end-user site.

| Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-06 |  |  |  |
| --- | --- | --- | --- |
| General Equipments |  |  |  |
| Sl.No. | DESCRIPTION OF HOSPITAL EQUIPMENT | Qty | EMD (Rs: |
| 1 | Table | 71 |  |
| 2 | Table & Chair | 3 |  |
| 3 | Refrigerator, hot zone, 210L | 6 |  |
| 4 | Wall Clock with second hand | 16 |  |
| 5 | Single Door Refrigerator (Domestic with Temp. Display) | 11 |  |
| 6 | Freezer | 12 |  |
| 7 | Ice Box | 12 |  |
| 8 | AC (1.5 Tonne) | 13 |  |

1). Table

Technical Specifications:

- Size: 2.5 Ft length x 2 Ft width x 3 Ft Height with one drawer

- 4 legs fitted with PVC rubber/stumps

- Table frame work made of Square M.S. tube, and Powder coated body with plywood top

- Should have smooth edges and robust construction body frame

- Table top should be washable top

Supplied with:

- 4 x PVC rubber/stumps

2). Table 3 Feet x 2 Feet with Drawer, and revolving stool with back rest

Technical Specifications:

- Size: 3 Ft length x 2 Ft width x 3 Ft Height with one drawer in right side

- 4 legs fitted with PVC Rubber/stumps

- Table frame work made of Square M.S. tube, and Powder coated body with plywood top

- Should have smooth edges and robust construction body frame

- Table top should be washable top

- Revolving stool with height adjustable with back rest

- Cushioned seat

- 5 legs base on caster for best load bearing

Supplied with:

- 4 x PVC Rubber/stumps

a). Revolving stool for Nurses station

Technical Specifications:

- It should be made of best quality of SS or MS frame, with foam with

- Should be round  in shape

- Should be light weight

- It should be robustness in quality

Supplied with:

- Supplied with foam rubber cushion in bottom legs

- Installation and training at end-user site

3) Refrigerator, Hot zone 210L

Technical Specifications:

- It should be reputed brand and make refrigerator

- Capacity should be Approx. 210 liters

- Should have a freezer compartment

- Should have automatic defrosting facility

- Should have five star electrical energy rating

- System accessories, spares & consumable system as specified

- Power supply, power input to be 220-240 VAC, 50 Hz

- Supplied with stand and stabilizer

- User manual with trouble shooting guidance, in English.

4). Wall Clock with Second hand

Technical Specifications:

- Should be reputed brand and make

- Supplied with second hand

- With one battery

- Round or squire in shape

- Light weight.

5). Single Door Refrigerator (Domestic with external Temp. Display)

Technical Specifications:

- Should be reputed make double door refrigerator

- Capacity should be Approx. 210 litres

- Should have a freezer compartment

- Should have automatic defrosting facility

- Should have five star electrical energy rating

- System accessories, Spares & Consumable

- System as specified

- Power Supply, power input to be 220-240 VAC, 50 Hz

- Supplied with stand and Stabilizer

- User manual with trouble shooting guidance, in English

6. Freezer - 20°C

Technical Specifications:

- Approximated capacity of 120 liters

- Should be constructed of double walls, the exterior is made of sheet steel while the interior of having PUF insulation, for minimize heat loss, powder painted body for glossy finish

- Should be low temperature up to - 20°C,

- Inner lining of Aluminum textured sheet ensuring

- Contemporary and sturdy design

- Corrosion resistant with outer body in powder-coat

- Insulation CFC - Free light density poly therein foam insulation

- Closure: Durable magnetic rubber gaskets for leak proof closure

- Main door: Elegant door trims, spring loaded hinges in doors, with sunken handles

- Heater Line: Hot air over the front door to prevent moisture condensation

- Refrigeration: hermetically sealed compressor with non CFC. HCPC eco friendly

- Refrigerant (R- 134A)

- Over load protection compressor with internal /external over load protector.

- Probe: Fast response PT-100 RTD sensor for accurate temperature measurement

- Safety: Slow blow Fuse & OLP’s

- Door lock for restricted accessibility

- Smooth moving caster wheels

- It should have minimum 3 star rating.

- It should work on 210/270 volts 50/60Hz, Single phase

- Device is produced by ISO 9001 certified manufacturer.

- Device is safety certified according CE or equivalent.

- Should have certificate of National Physical Laboratories (NPL).

- Accessories: Supplied with Powder card, 3pin moulded plug, Spare fuses.

7). Ice box Portable

Technical Specifications:

- It should be fully insulated.

- Size is to accommodate 25 to 50 blood bags.

- Same temperature should be maintained for longer time.

- It should be non breakable with good handle for carrying

- It should be ISI Marked box

8). Air Conditioner (1.5 Tone)

Technical Specifications:

- Should be reputed brand and make

- Supplied AC will be Split with digital display of temperature in indoor unit with star rating

- Supplied with outdoor stand, Stabilizer, with water pipe outlet standard accessories

- As per the market technical specification which is best in the market.

- 1 x set of connection tubing water drainage

- User manual with trouble shooting guidance, in English

- Training and installation at end-user site

SECTION VI

QUALIFICATION CRITERIA

(Referred to in Clause 1 of ITB)

(a)	The bidder or the manufacturer whose product is offered must have manufactured and supplied similar equipment of the type specified in the Bid Documents at least 75% (rounding off to nearest whole number) of the quantity offered in any one of the last  calendar years and should be satisfactorily functioning with no adverse report for at least one year on the date of bid opening.

(b) 	The bidder should furnish the information on past supplies and satisfactory performance in the Proforma given under Section - VIII.

(c) 	Bidders shall invariably furnish documentary evidence (Client’s Certificate) in support of the satisfactory operation of the equipment as specified above.

(d)	The Bidder shall furnish data to support that he has the financial and production capacity to perform the contract and complete the supplies within the stipulated delivery period.

SECTION VII

PROFORMA FOR EQUIPMENT AND QUALITY CONTROL EMPLOYED BY THE MANUFACTURER

(Please attach detailed proforma incorporating the information given below)

BID NO……………………….

DATE OF OPENING……..

NAME OF THE BIDDER:……………………………………..

(Note: All details should relate to the manufacturer for the items offered for supply).

1.	Name & Full address of the Manufacturer

(a) PAN No              (b). ECC Code                    ( c) IEC if applicable

(d) CST No               (e) TIN No                           (f)   Any other details

2.	(a) Telephone No.  				Office/Factory/Works

(b) Fax No.					Office/Factory/Works

(c) e-mail address

- Location of the manufacturing factory

- Details of important Plant & Machinery functioning in each dept..

- Details of the process of manufacturer in the factory

- Details & stocks of raw materials held

- Production capacity of item(s) quoted for, with the existing Plant &  Machinery

7.1 Normal                                  	7.2 Maximum

- Details of arrangement for quality control of products such as laboratory, testing equipment etc.

- Details of staff:

9.1 Details of technical supervisory staff in charge of production & quality control

  - Production: Managers/engineers/ officers/supervisors

  - QA / Q.C  :Managers/engineers/ officers/supervisors

  - Skilled labor employed

  - Unskilled labor employed

  - Maximum No. of workers (skilled & unskilled) employed on any day during the 18 months preceding the date of Tender.

- Whether Goods are tested to any standard specification?  If so, copies of specifications & original test certificates should be submitted.

- Whether bidder is willing to keep minimum stock of the item? If yes, how much can be kept?

- What is the lead-time for execution of order for mentioned quantity of the item?

SECTION VIII (a)

PROFORMA FOR PERFORMANCE STATEMENT

Proforma for Performance Statement (for a period of last 3 years)

Bid No.IFB NO: HLL/AFT/TDG/HP-II/EQUIP/2014-15/IV

Name of the firm:

| Order placed by (Full Address of Purchaser) | Order No. and Date | Description Of ordered item | Quantity of ordered item | Value of order | Date of Completion of delivery | Remarks indicating reasons for late delivery, if any | Has the equipment been satisfactorily functioning? (Attach a certificate from the Purchaser/Consignee) |
| --- | --- | --- | --- | --- | --- | --- | --- |

Signature and Seal of the Bidder ...................................

........................................................................................

SECTION VIII (b)

## HLL LIFECARE LIMITED

Akkulam Factory

THIRUVANANTHAPURAM-695017

IFB NO: HLL/AFT/TDG/HP-II/EQUIP/2014-15/IV

INDEMINITY CLAUSE

If the supplier fails to execute the order within the time prescribed for the delivery of goods ordered or violates or infringes the existing rates as agreed to as mentioned in the supply order, the supplier shall and will indemnify the company against all loses or damages whatsoever to be incurred or sustained including the legal cost or expenses incurred by the company by reason of non-delivery of goods at agreed quantity and rate with in the time specified in the supply order.  The company will initiate legal action if the supplier fails to execute the supply order as per the schedule in the supply order for the actual loss suffered. No quantity tolerance will be permitted in this regard unless otherwise prior approval is taken by the company before dispatching any excess quantity supplied which shall be returned back on freight to pay basis at the risk of the supplier.   Responsiveness of the Bid shall be at the discretion of HLL.

The supplier shall have no right to change the quantity stipulated in the supply order.

Bid pronounced Non Responsive by HLL shall be summarily rejected.

The decision of HLL will be final and no correspondence of this shall be entertained.

We have read and understood the above conditions and agree to abide   by the same.

PLACE:                                NAME AND SIGNATURE OF THE APPLICANT

DATE:                                                          (WITH OFFICE SEAL)

SECTION VIII (c)

## HLL LIFECARE LIMITED

Akkulam Factory

THIRUVANANTHAPURAM-695017

IFB NO: HLL/AFT/TDG/HP-II/EQUIP/2014-15/IV

#### DECLARATION

I / WE, The Bidder undertake, we shall execute, with in the purview of the contract, all the works and activities includes; transportation, loading, unloading and other technical work for the installation of the equipment at the designated site as per the contract / purchase order, at Six District wise Hospitals in Himachal Pradesh State.

We confirm having read and understood all the specifications, instructions, forms, terms and conditions and other requirements of the above tender (both expressed and implied) in full and that we agree to abide by all without any deviation.

SEAL OF THE APPLICANT

SIGNATURE

NAME AND ADDRESS OF APPLICANT

PART  B

SECTION IX

PRICE BID FORM

Date :..............................

IFB NO: HLL/AFT/TDG/HP-II/EQUIP/2014-15/IV

To:

HLL Lifecare Limited,

Akkulam Factory,

Sreekariam P.O.

Thiruvananthapuram - 17

Dear Sir

Having examined the Bidding Documents including Addenda Nos......................, the receipt of which is hereby duly acknowledged, we, the undersigned, offer to supply and deliver...........................................

(Description of Goods and Services) in conformity with the said Bidding Documents for the sum of.............................. (Total Bid amount in Words and Figures) or such other sums as may be ascertained in accordance with the Schedule of Prices attached herewith and made part of this bid.

We agree to abide by price per unit mentioned in price schedule, for the bid validity period specified in the bid document and it shall remain binding upon us and may be accepted by you at any time before the expiration of that period.

Until a formal final contract is prepared and executed between us, this bid, together with your written acceptance of the bid and your notification of award, shall constitute a binding Contract between us. We understand that you are not bound to accept the lowest or any bid you may receive.

We confirm that we comply with the eligibility requirements as per ITB Clause 1 of the bidding documents.

Dated this ....... day of ............................ 20 ......

___________________________________

(signature)

(in the capacity of) ____________________

Duly authorized to sign Bid for and on behalf of  ____________________________________

SECTION X

PRICE SCHEDULE (I)

Neonate Critical Equipments- Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-01

| S.No. | Item Description | Brand/Model | Manufacturer | Ex-factory price (Basic Unit Price exclusive of all levies & charges) | Excise Duty |  | Sales Tax |  | Insurance & Freight FOR Delivery destination (Himachal Pradesh |  | Unit Price inclusive of all levies & charges (ColumnNo. e+g+i+k) in Rs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | % | Amt. | % | Amt. | % | Amt. |  |
| a | B | c | d | e | f | g | h | I | j | k | l |
| 1. | Open care system: Radiant warmer, fixed height, with trolley, drawers, stand infusion double hook two poles attached. |  |  |  |  |  |  |  |  |  |  |
| 2 | Phototherapy unit, Double Surface, tachable , LED Type |  |  |  |  |  |  |  |  |  |  |

Delivery schedule : strictly with in 45 days (Third Week of July, 2014)

NOTE:  1) The columns may be filled as is applicable in each case.

2) Bidders should bid for the entire quantity asked by the purchaser in the tender document.

Bidding for part quantity for any equipment will not be considered.

*** 3) Product catalogue  along WITH COMPLIANCE SHEET confirming to technical specification should be attached with offer.

*** 4) List of optional accessories/ spares with rate, to be attached to price bid as separate sheet

Seal and Signature of the bidder

SECTION X

PRICE SCHEDULE (I)

Newborn Unit Equipments - Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-02

| S.No. | Item Description | Brand/Model | Manufacturer | Ex-factory price (Basic Unit Price exclusive of all levies & charges) | Excise Duty |  | Sales Tax |  | Insurance & Freight FOR Delivery destination (Himachal Pradesh |  | Unit Price inclusive of all levies & charges (ColumnNo. e+g+i+k) in Rs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | % | Amt. | % | Amt. | % | Amt. |  |
| a | B | c | d | e | f | g | h | I | j | k | l |
| 1 | Resuscitator, hand-operated, neonate, 250 ml |  |  |  |  |  |  |  |  |  |  |
| 2 | Resuscitator, hand-operated, neonate, 500ml |  |  |  |  |  |  |  |  |  |  |
| 3 | Oxygen hood S and M set of 3 each, including connecting tubes |  |  |  |  |  |  |  |  |  |  |
| 4 | Moniter, Vital sign, NIBP, HR, SpO2, ECG, RR and Temp. |  |  |  |  |  |  |  |  |  |  |
| 5 | Pulse oximeter, bedside, neonatal |  |  |  |  |  |  |  |  |  |  |
| 6 | Syringe pump, 10,20,50 ml, single phase |  |  |  |  |  |  |  |  |  |  |
| 7 | ECG unit, 3 channal, portable/SET |  |  |  |  |  |  |  |  |  |  |
| 8 | Oxygen concentrator |  |  |  |  |  |  |  |  |  |  |
| 9 | Pump, suction, Foot operated |  |  |  |  |  |  |  |  |  |  |
| 10 | Brest Pump |  |  |  |  |  |  |  |  |  |  |
| 11 | Laryngoscope set, neonate |  |  |  |  |  |  |  |  |  |  |
| 12 | Centrifuge, hematocrit, bench top, up to 4000 RPM, including rotor |  |  |  |  |  |  |  |  |  |  |
| 13 | Pump, suction, portable, 220V, w/access. |  |  |  |  |  |  |  |  |  |  |
| 14 | Stadio Meter |  |  |  |  |  |  |  |  |  |  |
| 15 | ABG Analysis Machine |  |  |  |  |  |  |  |  |  |  |
| 16 | Portable Ultra sound |  |  |  |  |  |  |  |  |  |  |
| 17 | Portable X - Ray machine |  |  |  |  |  |  |  |  |  |  |
| 18 | Flux meter |  |  |  |  |  |  |  |  |  |  |
| 19 | Nebulization Machine |  |  |  |  |  |  |  |  |  |  |
| 20 | CPAP |  |  |  |  |  |  |  |  |  |  |
| 21 | Fumigation Machine |  |  |  |  |  |  |  |  |  |  |

Delivery schedule : strictly with in 45 days (Third Week of July, 2014)

NOTE:  1) The columns may be filled as is applicable in each case.

2) Bidders should bid for the entire quantity asked by the purchaser in the tender document.

Bidding for part quantity for any equipment will not be considered.

*** 3) Product catalogue  along WITH COMPLIANCE SHEET confirming to technical specification should be attached with offer.

*** 4) List of optional accessories/ spares with rate, to be attached to price bid as separate sheet

Seal and Signature of the bidder

SECTION X

PRICE SCHEDULE (I)

Blood Bank Equipments - Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-03

| S.No. | Item Description | Brand/Model | Manufacturer | Ex-factory price (Basic Unit Price exclusive of all levies & charges) | Excise Duty |  | Sales Tax |  | Insurance & Freight FOR Delivery destination (Himachal Pradesh |  | Unit Price inclusive of all levies & charges (ColumnNo. e+g+i+k) in Rs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | % | Amt. | % | Amt. | % | Amt. |  |
| a | B | c | d | e | f | g | h | I | j | k | l |
| 1 | Blood Bank Refrigerator (2 o C - 8 o C) |  |  |  |  |  |  |  |  |  |  |
| 2 | Water Bath |  |  |  |  |  |  |  |  |  |  |
| 3 | Precision Incubator |  |  |  |  |  |  |  |  |  |  |
| 4 | Binocular Microscope |  |  |  |  |  |  |  |  |  |  |
| 5 | Table Top Centrifuge |  |  |  |  |  |  |  |  |  |  |

Delivery schedule : strictly with in 45 days (Third Week of July, 2014)

NOTE:  1) The columns may be filled as is applicable in each case.

2) Bidders should bid for the entire quantity asked by the purchaser in the tender document.

Bidding for part quantity for any equipment will not be considered.

*** 3) Product catalogue  along WITH COMPLIANCE SHEET confirming to technical specification should be attached with offer.

*** 4) List of optional accessories/ spares with rate, to be attached to price bid as separate sheet

Seal and Signature of the bidder

SECTION X

PRICE SCHEDULE (I)

Diesel Generator Sets - 5KVA & 25 KVA - Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-04

| S.No. | Item Description | Brand/Model | Manufacturer | Ex-factory price (Basic Unit Price exclusive of all levies & charges) | Excise Duty |  | Sales Tax |  | Insurance & Freight FOR Delivery destination (Himachal Pradesh |  | Unit Price inclusive of all levies & charges (ColumnNo. e+g+i+k) in Rs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | % | Amt. | % | Amt. | % | Amt. |  |
| a | B | c | d | e | f | g | h | I | j | k | l |
| 1 | Generator set 25 KVA |  |  |  |  |  |  |  |  |  |  |
| 2 | Generator 5 KVA |  |  |  |  |  |  |  |  |  |  |

Delivery schedule : strictly with in 45 days (Third Week of July, 2014)

NOTE:  1) The columns may be filled as is applicable in each case.

2) Bidders should bid for the entire quantity asked by the purchaser in the tender document.

Bidding for part quantity for any equipment will not be considered.

*** 3) Product catalogue  along WITH COMPLIANCE SHEET confirming to technical specification should be attached with offer.

*** 4) List of optional accessories/ spares with rate, to be attached to price bid as separate sheet

Seal and Signature of the bidder

SECTION X

PRICE SCHEDULE (I)

Hospital Equipments - Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-05

| S.No. | Item Description | Brand/Model | Manufacturer | Ex-factory price (Basic Unit Price exclusive of all levies & charges) | Excise Duty |  | Sales Tax |  | Insurance & Freight FOR Delivery destination (Himachal Pradesh |  | Unit Price inclusive of all levies & charges (ColumnNo. e+g+i+k) in Rs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | % | Amt. | % | Amt. | % | Amt. |  |
| a | B | c | d | e | f | g | h | I | j | k | l |
| 1 | Infant meter, plexus, 3½ft/105cm |  |  |  |  |  |  |  |  |  |  |
| 2 | Light, examination, mobile,220-12V |  |  |  |  |  |  |  |  |  |  |
| 3 | Spot Lamps |  |  |  |  |  |  |  |  |  |  |
| 4 | Crash Cart |  |  |  |  |  |  |  |  |  |  |
| 5 | Medicine Trolly |  |  |  |  |  |  |  |  |  |  |
| 6 | Basin, kidney, stainless steel, 825ml |  |  |  |  |  |  |  |  |  |  |
| 7 | Tray,dressing,ss,300x200x30mm |  |  |  |  |  |  |  |  |  |  |
| 8 | Glucometer with Dextrostix |  |  |  |  |  |  |  |  |  |  |
| 9 | X-Ray Viewing Box |  |  |  |  |  |  |  |  |  |  |
| 10 | Hub Cutter, syringe |  |  |  |  |  |  |  |  |  |  |
| 11 | E.Weighing Scale Digital - 100 Kg. |  |  |  |  |  |  |  |  |  |  |
| 12 | Scale, Baby, Electronic, 10 kg <5kg> |  |  |  |  |  |  |  |  |  |  |
| 13 | Weighing Scale Digital - 10Kg (Resolution up to 2 grams) |  |  |  |  |  |  |  |  |  |  |

Delivery schedule : strictly with in 45 days (Third Week of July, 2014)

NOTE:  1) The columns may be filled as is applicable in each case.

2) Bidders should bid for the entire quantity asked by the purchaser in the tender document.

Bidding for part quantity for any equipment will not be considered.

*** 3) Product catalogue  along WITH COMPLIANCE SHEET confirming to technical specification should be attached with offer.

*** 4) List of optional accessories/ spares with rate, to be attached to price bid as separate sheet

Seal and Signature of the bidder

SECTION X

PRICE SCHEDULE (I)

General Equipments - Tender No. IFB NO: HLL/AFT/TDG/HP-II/Equip/2014-15/IV-06

| S.No. | Item Description | Brand/Model | Manufacturer | Ex-factory price (Basic Unit Price exclusive of all levies & charges) | Excise Duty |  | Sales Tax |  | Insurance & Freight FOR Delivery destination (Himachal Pradesh |  | Unit Price inclusive of all levies & charges (ColumnNo. e+g+i+k) in Rs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | % | Amt. | % | Amt. | % | Amt. |  |
| a | B | c | d | e | f | g | h | I | j | k | l |
| 1 | Table |  |  |  |  |  |  |  |  |  |  |
| 2 | Table & Chair |  |  |  |  |  |  |  |  |  |  |
| 3 | Refrigerator, hot zone, 210L |  |  |  |  |  |  |  |  |  |  |
| 4 | Wall Clock with second hand |  |  |  |  |  |  |  |  |  |  |
| 5 | Single Door Refrigerator (Domestic with Temp. Display) |  |  |  |  |  |  |  |  |  |  |
| 6 | Freezer |  |  |  |  |  |  |  |  |  |  |

Delivery schedule : strictly with in 45 days (Third Week of July, 2014)

NOTE:  1) The columns may be filled as is applicable in each case.

2) Bidders should bid for the entire quantity asked by the purchaser in the tender document.

Bidding for part quantity for any equipment will not be considered.

*** 3) Product catalogue  along WITH COMPLIANCE SHEET confirming to technical specification should be attached with offer.

*** 4) List of optional accessories/ spares with rate, to be attached to price bid as separate sheet

Seal and Signature of the bidder

SECTION XI

Item - by - Item CMC Price Schedule (I)

| S.No. | Item Description | Brand/Model | Manufacturer | Year 1 Price |  |  | Year 2 Price |  |  | Year 3 Price |  |  | Year 4 Price |  |  | Year 5 Price |  |  | Grand Total Price |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Spares | Others | Total | Spares | Others | Total | Spares | Others | Total | Spares | Others | Total | Spares | Others | Total |  |
| a | B | d | E | F | G | I | j | k | l | m | n | o |  |  |  |  |  |  | p |

Scope of Contract:No of Visit (during the contract period)    a) Preventive maintenance visit:…………  b) Break down maintenance  visit:……..

“CMC” Comprehensive maintenanceContract includes;labors, spare and preventive & break down maintenance.

Note: Bidders should offer AMC for the entire quantity asked by the purchaser in the tender document.

Part AMC for part quantity for any equipment will not be considered.

Seal and Signature of the bidder

SECTION – XIII

BANK GUARANTEE FORM FOR EMD

Whereas ___________________ (hereinafter called the “Tenderer”) has submitted its quotation dated ___________________ for the supply of ____________________ (hereinafter called the “tender”) against the purchaser‟s tender enquiry No. __________________ Know all persons by these presents that we _________________________ of ______________________ (Hereinafter called the “Bank”) having our registered office at ________________________________________ are bound unto ____________________________ (hereinafter called the “Purchaser) in the sum of _________________________ for which payment will and truly to be made to the said Purchaser, the Bank binds itself, its successors and assigns by these presents.

Sealed with the Common Seal of the said Bank this ________day of _______ 20____. The conditions of this obligation are:

1) If the Tenderer withdraws or amends, impairs or derogates from the tender in any respect within the period of validity of this tender.

2) If the Tenderer having been notified of the acceptance of his tender by the Purchaser during the period of its validity:-

fails or refuses to furnish the performance security for the due performance of the contract or fails or refuses to accept/execute the contract or

if it comes to notice that the information/documents furnished in its tender is incorrect, false, misleading or forged

We undertake to pay the Purchaser up to the above amount upon receipt of its first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser will note that the amount claimed by it is due to it owing to the occurrence of one or both the two conditions, specifying the occurred condition(s). This guarantee will remain in force for a period of forty-five days after the period of tender validity and any demand in respect thereof should reach the Bank not later than the above date. …………………………….

(Signature with date of the authorised officer of the Bank)

………………………………………………………….

Name and designation of the officer

………………………………………………………….

………………………………………………………….

Seal, name & address of the Bank and address of the Branch

SECTION – XIV

MANUFACTURER‟S AUTHORISATION FORM

Joint General Manager (Mtls)

HLL LIFECARE LIMITED

AKKULAM  FACTORY,

SREEKARIAM  P.O.

THIRUVANANTHAPURAM - 17

Dear Sir,

Ref: Your TE document No ____________ dated _____________

We, ____________________________ who are proven and reputable manufacturers of___________________________(name and description of the goods offered in the tender) having factories at_____________________, hereby authorise Messrs________________(name and address of the agent) to submit a tender, process the same further and enter into a contract with you against your requirement as contained in the above referred TE documents for the above goods manufactured by us. We also state that we are not participating directly in this tender for the following reason(s): ___________________________________________________________(please provide reason here).

We further confirm that no supplier or firm or individual other than Messrs. __________________________ (name and address of the above agent) is authorised to submit a tender, process the same further and enter into a contract with you against your requirement as contained in the above referred TE documents for the above goods manufactured by us. We also hereby extend our full warranty, CMC as applicable as per clause 15 of the General Conditions of Contract, read with modification, if any, in the Special Conditions of Contract for the goods and services offered for supply by the above firm against this TE document. We also hereby confirm that we would be responsible for the satisfactory execution of contract placed on the authorised agent We also confirm that the price quoted by our agent shall not exceed the price which we would have quoted directly”

Yours faithfully,

[Signature with date, name and designation]

for and on behalf of Messrs___________________________

[Name & address of the manufacturers]

Note: 1. This letter of authorisation should be on the letter head of the manufacturing firm and should be signed by a person competent and having the power of attorney to legally bind the manufacturer.

2.Orginal Letter may be sent. Photocopy not acceptable.

SECTION – XV

BANK GUARANTEE FORM FOR PERFORMANCE SECURITY/ CMC SECURITY

The Joint General Manager (Mtls)

HLL LIFECARE LIMITED

AKKULAM  FACTORY,

SREEKARIAM  P.O.

THIRUVANANTHAPURAM - 17

WHEREAS _____________________________ (Name and address of the supplier) (Hereinafter called “the supplier”) has undertaken, in pursuance of contract no________________________ dated _____________ to supply (description of goods and services) (herein after called “the contract”). AND WHEREAS it has been stipulated by you in the said contract that the supplier shall furnish you with a bank guarantee by a scheduled commercial bank recognised by you for the sum specified therein as security for compliance with its obligations in accordance with the contract;

AND WHEREAS we have agreed to give the supplier such a bank guarantee;

NOW THEREFORE we hereby affirm that we are guarantors and responsible to you, on behalf of the supplier, up to a total of. __________________ (Amount of the guarantee in words and figures), and we undertake to pay you, upon your first written demand declaring the supplier to be in default under the contract and without cavil or argument, any sum or sums within the limits of (amount of guarantee) as aforesaid, without your needing to prove or to show grounds or reasons for your demand or the sum specified therein.

We hereby waive the necessity of your demanding the said debt from the supplier before presenting us with the demand.

We further agree that no change or addition to or other modification of the terms of the contract to be performed there under or of any of the contract documents which may be made between you and the supplier shall in any way release us from any liability under this guarantee and we hereby waive notice of any such change, addition or modification. This guarantee shall be valid up to 30 (thirty) months from the date of Notification of Award i.e. up to ----------- (indicate date) …………………………….

(Signature with date of the authorised officer of the Bank)

………………………………………………………….

Name and designation of the officer

………………………………………………

…………. …………………………

……………………………….

Seal, name & address of the Bank and address of the Branch

SECTION – XVI

CONTRACT FORM - A

CONTRACT FORM FOR SUPPLY, INSTALLATION, COMMISSIONING, HANDING OVER, TRIAL RUN, TRAINING OF OPERATORS & WARRANTY OF GOODS

……………………………..(Address of the Purchaser‟s/Consignee‟s office issuing the contract) Contract No___________ dated______________ This is in continuation to this office‟s Notification of Award No_______ dated ______ 1. Name & address of the Supplier: ______________________________ 2. Purchaser‟s TE document No________ dated____________ and subsequent Amendment No____________, dated_________ (if any), issued by the purchaser 3. Supplier‟s Tender No_________ dated__________ and subsequent communication(s) No____________ dated _________ (if any), exchanged between the supplier and the purchaser in connection with this tender. 4. In addition to this Contract Form, the following documents etc, which are included in the documents mentioned under paragraphs 2 and 3 above, shall also be deemed to form and be read and construed as integral part of this contract:

(i) General Conditions of Contract;

(ii) Special Conditions of Contract;

(iii) List of Requirements;

(iv) Technical Specifications;

(v) Quality Control Requirements;

(vi) Tender Form furnished by the supplier;

(vii) Price Schedule(s) furnished by the supplier in its tender;

(viii) Manufacturers‟ Authorisation Form (if applicable for this tender);

(ix) Purchaser‟s Notification of Award

Note: The words and expressions used in this contract shall have the same meanings as are respectively assigned to them in the conditions of contract referred to above. Further, the definitions and abbreviations incorporated under General Instructions to Tenderers‟ of the Purchaser‟s TE document shall also apply to this contract, Some terms, conditions, stipulations etc. out of the above-referred documents are reproduced below for ready reference:

(i) Brief particulars of the goods and services which shall be supplied/ provided by the supplier are as under:

a)…………………………….

b)………………………………

c)…………………………….. Any other additional services (if applicable) and cost thereof: ___________________ Total value (in figure) ____________ (In words) ___________________________

(ii) Delivery schedule

(iii) Details of Performance Security

(iv) Quality Control

(a) Mode(s), stage(s) and place(s) of conducting inspections and tests.

(b) Designation and address of purchaser‟s inspecting officer

(v) Destination and despatch instructions

(vi) Consignee, including port consignee, if any

6. Warranty clause

7. Payment terms

8. Paying authority

(Signature, name and address of the Purchaser‟s/Consignee‟s authorised official)

For and on behalf of__________________

Received and accepted this contract

(Signature, name and address of the supplier‟s executive duly authorised to sign on behalf of the supplier)

For and on behalf of _________________________

(Name and address of the supplier)

(Seal of the supplier)

Date: _________________________ Place: _________________________

SECTION – XVII

CONSIGNEE RECEIPT CERTIFICATE

(To be given by consignee‟s authorized representative)

The following store (s) has/have been received in good condition:

1) Contract No. & date :______________________________

2) Supplier‟s Name :______________________________

3) Consignee‟s Name & Address with telephone No. & Fax No. : ______________________________

4) Name of the item supplied :______________________________

5) Quantity Supplied :______________________________

6) Date of Receipt by the Consignee :______________________________

7) Name and designation of Authorized Representative of Consignee :______________________________

8) Signature of Authorized Representative of Consignee with date :______________________________

9) Seal of the Consignee :______________________________

SECTION – XVIII

Proforma of Final Acceptance Certificate by the Consignee

No_______________ Date_______________

To

M/s _______________________

_______________________

_______________________

Subject: Certificate of commissioning of equipment/plant.

This is to certify that the equipment(s)/plant(s) as detailed below has/have been received in good conditions along with all the standard and special accessories and a set of spares (subject to remarks in Para no.02) in accordance with the contract/technical specifications. The same has been installed and commissioned.

(a) Contract No______________________________________ dated_______________

(b) Description of the equipment(s)/plants: ____________________________________

(c) Equipment(s)/ plant(s) nos.:_____________________________________________

(d) Quantity: ____________________________________________________________

(e) Name of the vessel/Transporters:__________________________________________

(f) Name of the Consignee:_________________________________________________

(g) Date of commissioning and proving test:____________________________________

Details of accessories/spares not yet supplied and recoveries to be made on that account.:

| Sl. No. | Description of Item | Quantity | Amount to be recovered |
| --- | --- | --- | --- |

The proving test has been done to our entire satisfaction and operators have been trained to operate the equipment(s)/plant(s).

The supplier has fulfilled its contractual obligations satisfactorily ## or

The supplier has failed to fulfil its contractual obligations with regard to the following:

a) He has not adhered to the time schedule specified in the contract in dispatching the documents/ drawings pursuant to „Technical Specifications‟.

b) He has not supervised the commissioning of the equipment(s)/plant(s)in time, i.e. within the period specified in the contract from date of intimation by the Purchaser/Consignee in respect of the installation of the equipment(s)/plant(s).

c) The supplier as specified in the contract has not done training of personnel.

The extent of delay for each of the activities to be performed by the supplier in terms of the contract is The amount of recovery on account of non-supply of accessories and spares is given under Para no.02. The amount of recovery on account of failure of the supplier to meet his contractual obligations is_____________ (here indicate the amount).

(Signature) (Name)

(Designation with stamp)

## Explanatory notes for filling up the certificate:

i) He has adhered to the time schedule specified in the contract in dispatching the documents/drawings pursuant to „Technical Specification‟.

ii) He has supervised the commissioning of the equipment(s)/plant(s) in time, i.e. within the time specified in the contract from date of intimation by the Purchaser/Consignee in respect of the installation of the equipment(s)/plant(s).

iii) Training of personnel has been done by the supplier as specified in the contract.

iv) In the event of documents/drawings having not been supplied or installation and commissioning of the equipment(s)/plant(s) having been delayed on account of the supplier, the extent of delay should always be mentioned in clear terms

Section – XXI

Consignee List

SECTION XII

CHECK LIST

| Sl. No. | Documents | Page No. | Remarks |
| --- | --- | --- | --- |
| 1 | Bid Security in the form of DD drawn in favor of HLL Lifecare Limited, Payable at Thiruvananthapuram |  |  |
| 2 | EMD in the form of DD drawn in favor of HLL Lifecare Limited, Payable at Thiruvananthapuram / Bank Guarantee |  |  |
| 3 | Market standing certificate to prove that the bidder manufactures the quoted items for the past years |  |  |
| 4 | Duly attested copy of factory license/ Industrial license |  |  |
| 5 | Duly attested copy of sales tax registration certificate |  |  |
| 6 | Duly attested copy of document to prove the legal status, place of registration and principal place of business of the undertaking |  |  |
| 7 | Duly attested copies of quality certificates for the products and quality system certifications |  |  |
| 8 | Copy of Balance sheet for the past financial years, duly certified by a chartered accountant |  |  |
| 9 | Copy of Profit & Loss Account for years certified by chartered accountant |  |  |
| 10 | Duly filled proforma showing details of Equipment and Quality Control employed by the manufacturer as per section VII |  |  |
| 11 | Proforma for Performance Statement (for a period of last years) as per Section VIII |  |  |
| 12 | Self attested statement showing financial capability necessary to perform the Contract |  |  |
| 13 | Self attested statement showing technical and production capability necessary to perform the Contract as per Section VI |  |  |
| 14 | Client's Certificate as per Section VI |  |  |
| 15 | Product brochure literature, write up etc. |  |  |
| 16 | Item-by-item commentary (compliance statement) on the Purchaser’s Technical Specifications demonstrating substantial responsiveness of the goods and services to those specifications or a statement of deviations and exceptions to the provisions of the Technical Specifications |  |  |
| 17 | The letter of authorization indicated by written power-of-attorney |  |  |
| 18 | Spare parts and after sales service facilities in India |  |  |
| 19 | Documentary evidence of constitution of firm such as Memorandum of Articles, Partnership Deed, etc., with details of Name, Address, Tel. No., Fax No., E-mail Address of firm and the Managing Director / Partner / Proprietor |  |  |
| 20 | Annual turnover statement for last years certified by the Auditor |  |  |
| 21 | List of items quoted (without prices) |  |  |