## TENDER FOR SETTING UP OF MICROBIOLOGY SECTION ON TURNKEY BASIS & PROVIDING CONSULTANCY SERVICES FOR NABL ACCREDITATION

## FOR STATE DRUG TESTING LABORATORY, TRIPURA

### IFB NO: HLL/AFT-CMO/TPA /SDTL /2019-20, Dt. 08.07.2019

##### HLL LIFECARE LIMITED

(A Govt. of India Enterprise)

AKKULAM  FACTORY, SREEKARIAM  P.O.

THIRUVANANTHAPURAM – 17

KERALA, INDIA

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: : hcdcmo@lifecarehll.com

Web: www.lifecarehll.com

CIN: U25193KL1966GOI002621

INDEX

| Section | Topic | Page No. |
| --- | --- | --- |
| I | Notice inviting Tender (NIT) | 3 |
| II | General Instructions to Tenderers (GIT) | 7 |
| III | General Conditions of Contract (GCC) | 23 |
| IV | Special Conditions of Contract (SCC) | 37 |
| V | List of Requirements | 40 |
| VI | Technical Specifications | 43 |
| VII | Quality Control Requirements | 70 |
| VIII | Qualification Criteria | 71 |
| IX | Tender Form for Technical Bid | 72 |
| IX A | Proforma ‘A’ I & A II | 73 |
| IX B | Tender form for price bid | 75 |
| X | Price Schedules (A & B) | 76 |
| XI | Bank Guarantee Form for EMD | 110 |
| XII | Bank Guarantee Form for Performance Security /CAMC Security | 111 |
| XIII | Manufacturer’s/ Distributor’s Authorisation Form/ Consortium | 112 |
| XIV | Contract Form ’A’ | 116 |
| XV | Contract Form ‘B’ | 120 |
| XVI | Proforma of Consignee Receipt Certificate | 123 |
| XVII | Proforma of Final Acceptance Certificate by the Consignee | 124 |
| XVIII | Consignee Address | 126 |
| XIX | Check list | 127 |

HLL LIFECARE LIMITED

(A Govt. of India Enterprise)

AKKULAM  FACTORY, SREEKARIAM  P.O.

THIRUVANANTHAPURAM – 17, KERALA, INDIA

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: : hcdcmo@lifecarehll.com

Web: www.lifecarehll.com

### IFB NO: HLL/AFT-CMO/TPA/SDTL/2019-20,                                                           Dt. 08.07.2019

SECTION - I

NOTICE INVITING TENDERS

HLL Lifecare Limited Akkulam Factory, Sreekariam PO, Thiruvananthapuram – 17 hereby invites sealed bids under Two-Bid system from eligible suppliers/manufacturers/Consortium partners for setting up of Microbiology Section on turnkey basis and provide consultancy services for NABL accreditation for State Drug Testing Laboratory, Tripura.

Tender Documents will be issued from the Office of the Joint General Manager (Materials), HLL Lifecare Limited,  Akkulam Factory, Sreekariam PO, Thiruvananthapuram – 17, from 08.07.2019 to 29.07.2019 on all working days between 11.00 AM and 4.00 PM, on payment of  non-refundable fee of Tender cost of respective equipment (s). The tender may also be downloaded from our web site www.lifecarehll.com and the tender cost shall be submitted along with the tender in the form of Demand Draft taken in favour of HLL Lifecare Limited payable at Thiruvananthapuram. Failing to submit the tender document fee, the bid is liable to be rejected. SSI/MSME units interested in availing exemption from payment of Tender Fee & EMD should submit a valid copy of their registration certificate issued by the concerned DIC or NSIC. SSI/MSME units are not exempted from the Security Deposit, if the tender is awarded.

Last date and time of receipt of Tender 			: - 29.07.2019 at  14.00 PM.

Date and time of opening of Tender 				: - 29.07.2019 at  15.00 PM.

Deputy General Manager (Materials)

SCHEDULE FOR SUBMISSION OF APPLICATION

| EVENT | DATE |
| --- | --- |
| Starting date of sale of documents | 08.07.2019 |
| Last date of sale of documents | 29.07.2019 |
| Last date and time for submission of completed Tender | 29.07.2019 at 14.00 PM |
| Date and time for Opening of Technical Bid | 29.07.2019 at 15.00 PM |

The completed Tender should be submitted before the due date and time of submission at the following address.

DEPUTY GENERAL MANAGER (MATERIALS),

HLL LIFECARE LTD

AKKULAM  FACTORY, SREEKARIAM  P.O.

THIRUVANANTHAPURAM – 17, KERALA, INDIA

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: : hcdcmo@lifecarehll.com

Web: www.lifecarehll.com

SECTION - II

GENERAL INSTRUCTIONS TO TENDERERS (GIT)

CONTENTS

| Sl. No. | Topic | Page No. |
| --- | --- | --- |
| A | PREAMBLE |  |
| 1 | Definitions and Abbreviations | 7 |
| 2 | Introduction | 8 |
| 3 | Availability of Funds | 8 |
| 4 | Language of Tender | 8 |
| 5 | Eligible Tenderers | 9 |
| 6 | Eligible Goods and Services | 9 |
| 7 | Tendering Expense | 9 |
| B | TENDER ENQUIRY DOCUMENTS |  |
| 8 | Contents of Tender Enquiry Documents | 9 |
| 9 | Amendments to Tender Enquiry Documents | 10 |
| 10 | Clarification of Tender Enquiry Documents | 10 |
| C | PREPARATION OF TENDERS |  |
| 11 | Documents Comprising the Tender | 10 |
| 12 | Tender Currencies | 11 |
| 13 | Tender Prices | 11 |
| 14 | Indian Agent | 12 |
| 15 | Firm Price | 12 |
| 16 | Alternative Tenders | 12 |
| 17 | Documents Establishing Tenderer’s Eligibility and Qualifications | 12 |
| 18 | Documents Establishing Good’s Conformity to Tender Enquiry Document | 13 |
| 19 | Earnest Money Deposit (EMD) | 13 |
| 20 | Tender Validity | 14 |
| 21 | Signing and Sealing of Tender | 14 |

| D | SUBMISSION OF TENDERS |  |
| --- | --- | --- |
| 22 | Submission of Tenders | 15 |
| 23 | Late Tender | 15 |
| 24 | Alteration and Withdrawal of Tender | 16 |
| E | TENDER OPENING |  |
| 25 | Opening of Tenders | 16 |
| F | SCRUTINY AND EVALUATION OF TENDERS |  |
| 26 | Basic Principle | 16 |
| 27 | Responsiveness | 17 |
| 28 | Minor Infirmity/Irregularity/Non-Conformity | 18 |
| 29 | Discrepancy in Prices | 18 |
| 30 | Qualification Criteria | 18 |
| 31 | Tender currency( Indian Rupees) | 18 |
| 32 | Evaluation | 18 |
| 33 | Comparison of Tenders | 18 |
| 34 | Additional Factors and Parameters for Evaluation and Ranking of Responsive Tenders | 19 |
| 35 | Tenderer’s capability to perform the contract | 19 |
| 36 | Contacting the Purchaser | 19 |
| G | AWARD OF CONTRACT |  |
| 37 | Purchaser’s Right to Accept any Tender and to Reject any or All Tenders | 19 |
| 38 | Award Criteria | 19 |
| 39 | Variation of Quantities at the Time of Award | 20 |
| 40 | Notification of Award | 20 |
| 41 | Issue of Contract | 20 |
| 42 | Non-receipt of Performance Security and Contract by the Purchaser | 20 |
| 43 | Return of EMD | 20 |
| 44 | Corrupt or Fraudulent Practices | 20 |

SECTION – II

GENERAL INSTRUCTIONS TO TENDERERS (GIT)

A. PREAMBLE

1.	Definitions and Abbreviations

  - The following definitions and abbreviations, which have been used in these documents shall have the meanings as indicated below:

1.2.	Definitions:

(i)	“Purchaser” means the organization purchasing goods and services as incorporated in the Tender Enquiry document.

(ii)	“Tender” means Bids / Quotation / Tender received from a Firm / Tenderer /Bidder/Consortium.

(iii)	“Tenderer” means Bidder/ the Individual or Firm  or Consortium submitting Bids / Quotation / Tender

(iii)	“Supplier” means the individual or the firm supplying the goods and services as incorporated in the contract.

(iv)	“Goods” means the articles, material, commodities, furniture, fixtures, raw material, spares, instruments, machinery, equipment, medical equipment, etc. which the supplier is required to supply to the purchaser under the contract.

(v)	“Services” means services allied and incidental to the supply of goods, such as transportation, installation, commissioning, provision of technical assistance, training, after sales service, maintenance service and other such obligations of the supplier covered under the contract.

(vi)	“Earnest Money Deposit” (EMD) means Bid Security/ monetary or financial guarantee to be furnished by a tenderer along with its tender.

(vii)	“Contract” means the written agreement entered into between the purchaser and/or consignee and the supplier, together with all the documents mentioned therein and including all attachments, annexure etc. therein.

(viii)	“Performance Security” means monetary or financial guarantee to be furnished by the successful tenderer for due performance of the contract placed on it. Performance Security is also known as Security Deposit.

(ix)	“Consignee” means the Hospital/Institute  to whom the goods are required to be delivered as specified in the Contract.

(x)	“Specification” means the document/standard that prescribes the requirement with which goods or service has to conform.

(xi)	“Inspection” means activities such as measuring, examining, testing, gauging one or more characteristics of the product or service and comparing the same with the specified requirement to determine conformity.

(xii)	“Day” means calendar day.

1.3 	Abbreviations:

(i)	“T E Document” means Tender Enquiry Document

(ii)	“NIT” means Notice Inviting Tenders.

(iii)	“GIT” means General Instructions to Tenderers

(iv)	“SIT” means Special Instructions to Tenderers

(v)	“GCC” means General Conditions of Contract

(vi)	“SCC” means Special Conditions of Contract

(vii)	“DGS&D” means Directorate General of Supplies and Disposals

(viii)	“NSIC” means National Small Industries Corporation

(ix)	“PSU” means Public Sector Undertaking

(x)	“CPSU” means Central Public Sector Undertaking

(xi)	“SSI” means Small Scale Industry

(xii)	“LC” means Letter of Credit

(xiii)	“DP” means Delivery Period

(xiv)	“BG” means Bank Guarantee

(xv)	“GST” Goods and Service Tax

(xvi)       “RR” means Railway Receipt

(xvii)      “BL” means Bill of Lading

(xviii)     “FOB” means Free on Board

(xix)       “FCA” means Free Carrier

(xx)        “FOR” means Free On Rail

(xxi)       “CIF” means Cost, Insurance and Freight

(xxii)    “CIP (Destinations)” means Carriage and Insurance Paid up to named port of destination. Additionally the Insurance, local transportation and storage shall be extended and borne by the Supplier from ware house to the consignee site for a period including 3 months beyond date of delivery

(xxiii)     “DDP” means Delivery Duty Paid at named place of destination (consignee site)

(xxiv)    “INCOTERMS” means International Commercial Terms as on the date of Tender Opening

(xxv)      “RT” means Re-Tender

(xxvi)      “CAMC” Comprehensive Annual Maintenance Contract

2.	Introduction

2.1	The Purchaser has issued these TE documents for purchase of goods and related services as mentioned in Section – V and VI – “List of Requirements” and “Technical Specifications”.

2.2	This section (Section II - “General Instruction Tenderers”) provides the relevant information as well as instructions to assist the prospective tenderers in preparation and submission of tenders. It also includes the mode and procedure to be adopted by the purchaser for receipt and opening as well as scrutiny and evaluation of tenders and subsequent placement of contract.

2.3	Before formulating the tender and submitting the same to the purchaser, the tenderer should read and examine all the terms, conditions, instructions, checklist etc. contained in the TE documents. Failure to provide and/or comply with the required information, instructions etc. incorporated in these TE documents may result in rejection of its tender.

3.	Availability of funds

Deleted.

4.	Language of Tender

4.1 	The tender submitted by the tenderer and all subsequent correspondence and documents relating to the tender exchanged between the tenderer and the purchaser, shall be written in the English/Hindi language, unless otherwise specified in the Tender Enquiry. However, the language of any printed literature furnished by the tenderer in connection with its tender may be written in any other language provided the same is accompanied by an English translation and, for purposes of interpretation of the tender, the English translation shall prevail.

4.2	The tender submitted by the tenderer and all subsequent correspondence and documents relating to the tender exchanged between the tenderer and the purchaser, may also be written in the Hindi language, provided that the same are accompanied by English translation, in which case, for purpose of interpretation of the tender etc, the English translations shall prevail

5.	Eligible Tenderers

5.1 	This invitation for tenders is open to all suppliers who fulfil the eligibility criteria specified in these documents.

6.	Eligible Goods and Services

6.1	All goods and related services to be supplied under the contract shall have their origin in India or any other country with which India has not banned trade relations. The term “origin” used in this clause means the place where the goods are manufactured or from where the related services are arranged and supplied.

7.	Tendering Expense

7.1	The tenderer shall bear all costs and expenditure incurred and/or to be incurred by it in connection with its tender including preparation, mailing and submission of its tender and for subsequent processing the same. The purchaser will, in no case be responsible or liable for any such cost, expenditure etc regardless of the conduct or outcome of the tendering process.

B. TENDER DOCUMENTS

8.	Content of Tender Documents

8.1	In addition to Section I – “Notice inviting Tender” (NIT), the Tender documents include:

- Section II	– General Instructions to Tenderers (GIT)

- Section III	– General Conditions of Contract (GCC)

- Section IV	– Special Conditions of Contract (SCC)

- Section V	– List of Requirements

- Section VI	– Technical Specifications

- Section VII	– Quality Control Requirements

- Section VIII – Qualification Criteria

- Section IX	– Tender Form for Technical Bid

- Section IX B – Tender Form for Price Bid

- Section X	– Price Schedules

- Section XI	– Bank Guarantee Form for EMD

- Section XII – Bank Guarantee Form for Performance Security/CAMC Security

- Section XIII – Manufacturer’s/Distributor’s Authorisation Form

- Section XIV– Contract Form’ A ‘

- Section XV  – Contract Form ‘B’

- Section XVI – Proforma of Consignee Receipt Certificate

- Section XVII – Proforma of Final Acceptance Certificate by the consignee

- Section XVIII – Consignee List

- Section XIX - Check List

8.2 	The relevant details of the required goods and services, the terms, conditions and procedure for tendering, tender evaluation, placement of contract, the applicable contract terms and, also, the standard formats to be used for this purpose are incorporated in the above-mentioned documents. The interested tenderers are expected to examine all such details etc to proceed further.

9.	Amendments to Tender documents

9.1 	At any time prior to the deadline for submission of tenders, the purchaser may, for any reason deemed fit by it, modify the TE documents by issuing suitable amendment(s) to it.

9.2 	Such an amendment will be notified in Purchaser’s Website www.lifecarehll.com  and also in writing by registered/speed post or by e-mail, to all prospective tenderers, which have received the TE documents and will be binding on them.

9.3 	In order to provide reasonable time to the prospective tenderers to take necessary action in preparing their tenders as per the amendment, the purchaser may, at its discretion extend the deadline for the submission of tenders and other allied time frames, which are linked with that deadline.

10.	Clarification of Tender documents

10.1 	A tenderer requiring any clarification or elucidation on any issue of the TE documents may take up the same with the purchaser in writing. The purchaser will respond in writing to such request provided the same is received by the purchaser not later than 5 days prior to the prescribed date of submission of tender.

- C. PREPARATION OF TENDERS

11.	Documents Comprising the Tender

11.1	The Two bid System, i.e. “Techno–Commercial Tender” and “Price Tender” prepared by the tenderer shall comprise the following:

A)	Techno–Commercial Tender (Un priced Tender)

i)	Earnest money furnished in accordance with GIT clause 19.1

ii)	Tender Form as per Section IX (Un priced).

iii)	Documentary evidence, as necessary in terms of clauses GIT Clause 5 and 17 establishing that the tenderer is eligible to submit the tender and, also, qualified to perform the contract if its tender is accepted.

iv)	Tenderer/Agent who quotes for goods manufactured by other manufacturer shall furnish Manufacturer’s/ Distributor’s Authorisation Form/Consortium as per the format in Section XIII.

v)	Power of Attorney in favour of signatory of TE documents and signatory of Manufacturer’s/Distributor’s Authorisation Form.

vi)	Documents and relevant details to establish in accordance with GIT clause 18 that the goods and the allied services to be supplied by the tenderer conform to the requirement of the TE documents.

vii)	Performance Statement as per section XVII along with relevant copies of orders and end users’ satisfaction certificate.

viii)	Price Schedule(s) as per Section X filled up with all the details including Make, Model etc. of the goods offered with prices blank (without indicating any prices).

ix)	Certificate of Incorporation in the country of origin.

- Checklist as per Section XIX.

- Technical Compliance Statement (specification points-wise) along with pamphlets/Catalogue.

B)	Price Tender:

The information given at clause no. 11.1 A) ii) & viii) above should be reproduced with the prices indicated.

#### N.B.

- All pages of the Tender should be page numbered and indexed.

It is the responsibility of tenderer to go through the TE document to ensure furnishing

all required documents in addition to above, if any.

11.2 	The authorized signatory of the tenderer must sign the tender duly stamped at appropriate places and initial all the remaining pages of the tender.

11.3	A tender, which does not fulfil any of the above requirements and/or gives evasive information/reply against any such requirement, shall be liable to be ignored and rejected.

11.4   Tender sent by fax/telex/cable/electronically shall be ignored.

12.     Tender currencies

12.1 	The tenderer shall quote only in Indian Rupees.

Tenders, where prices are quoted in any other currency  shall be treated as non -responsive and rejected.

13	Tender Prices

13.1 	The Tenderer shall indicate on the Price Schedule provided under Section X all the specified components of prices shown therein including the unit prices and total tender prices of the goods and services. All the columns shown in the price schedule should be filled up as required. If any column doesn’t apply to a tender, same should be clarified as “NA” by the tenderer.

13.2 	Deleted.

13.3 	While filling up the columns of the Price Schedule, the following aspects should be noted for compliance:

a)	GST which will be payable on the goods shall be indicated in the SECTION X (a). in absence of GST, it will be treated as inclusive.

b)	Charges towards Packing & Forwarding, Inland Transportation, Insurance (local transportation and storage) would be borne by the supplier from ware house to the consignee site for a period including 3 months beyond date of delivery, Loading/Unloading, Positioning and other local costs incidental to delivery of the goods to their final destination as specified in Price Schedule;

c)	The price of Incidental Services, as mentioned in Section V & VI, Price Schedule and GCC shall be considered;

d)	The price of CAMC including all taxes and duties, as mentioned in GCC, SCC, List of Equipments, Technical Specification and Price Schedule shall be considered.

13.4	Additional information and instruction on GST:

13.4.1The price will be taken inclusive of GST and no claim for the same will be entertained later.

13.5	For transportation of imported goods offered from abroad, relevant instructions as incorporated under GCC Clause 10 shall be followed.

13.6	For insurance of goods to be supplied, relevant instructions as provided under GCC Clause 11 shall be followed.

13.7	The need for indication of all such price components by the tenderers, as required in this clause (viz., GIT clause 13) is for the purpose of comparison of the tenders by the purchaser and will no way restrict the purchaser’s right to award the contract on the selected tenderer on any of the terms offered.

14.	Indian Agent

14.1	If a foreign tenderer, the tenderer should have engaged an agent in India in connection with its tender and shall furnish the following information:

a)	The complete name, address, contact person and contact number of the Indian Agent(s)/service centres and its permanent income tax account number as allotted by the Indian Income Tax authority.

b)	The details of the services to be rendered by the agent for the subject requirement.

c)	Details of Service outlets in India, nearest to the consignee(s), to render services during Warranty and CAMC period.

15.	Firm Price

15.1	Unless otherwise specified, prices quoted by the tenderer shall remain firm and fixed during the currency of the contract and not subject to variation on any account.

15.2	However, as regards taxes and duties, if any, chargeable on the goods and payable, the conditions stipulated in GIT clause 13 will apply.

16.	Alternative Tenders

16.1	Alternative Tenders are not permitted.

17 	Documents Establishing Tenderer’s Eligibility and Qualifications

17.1	Pursuant to GIT clause 11, the tenderer shall furnish, as part of its tender, relevant details and documents establishing its eligibility to quote and its qualifications to perform the contract if its tender is accepted.

17.2	The documentary evidence needed to establish the tenderer’s qualifications shall fulfil the following requirements:

a)	In case the tenderer offers to supply goods, which are manufactured by some other firm, the tenderer has been duly authorised by the goods manufacturer to quote for and supply the goods to the purchaser. The tenderer shall submit the manufacturer’s/Distributor’s authorization letter /consortium to this effect as per the standard form provided under Section XIV in this document.

b)	The tenderer has the required financial, technical and production capability necessary to perform the contract and, further, it meets the qualification criteria incorporated in the Section VIII in these documents.

c)	In case the tenderer is not doing business in India, it is duly represented by an agent stationed in India fully equipped and able to carry out the required contractual functions and duties of the supplier including payment of  duties, taxes, levies, clearance of goods, freight, transport, insurance after sale service, maintenance & repair etc. of the goods in question, stocking of spare parts and fast moving components and other obligations, if any, specified in the conditions of contract and/or technical specifications.

d)	The tenderer shall quote for all items in the list of requirements, if any of the item not quoted then their bid will be treated as non responsive.

e)	The items quoted shall be responsive to the requirements in the List of Equipments in the Tender.

18.	Documents establishing Good’s Conformity to Tender document.

18.1 	The tenderer shall provide in its tender the required as well as the relevant documents like technical data, literature, drawings etc. to establish that the goods and services offered in the tender fully conform to the goods and services specified by the purchaser in the Tender documents. For this purpose the tenderer shall also provide a clause-by-clause commentary on the technical specifications and other technical details incorporated by the purchaser in the Tender documents to establish technical responsiveness of the goods and services offered in its tender.

18.2 	In case there is any variation and/or deviation between the goods & services prescribed by the purchaser and that offered by the tenderer, the tenderer shall list out the same in a chart form without ambiguity and provide the same along with its tender.

18.3	If a tenderer furnishes wrong and/or misguiding data, statement(s) etc. about technical acceptability of the goods and services offered by it, its tender will be liable to be ignored and rejected in addition to other remedies available to the purchaser in this regard.

19.	Earnest Money Deposit (EMD)

19.1

| Sl No | Description | Tender Fee In Rs. | EMD amount in Rs. |
| --- | --- | --- | --- |
| 1 | Setting up of Microbiology Section on Turnkey basis - Phase 1 | 5,000.00 | 2,50,000.00 |
| 2 | Providing Consultancy services for NABL accreditation for SDTL, Tripura – Phase 2 |  |  |

- Earnest Money shall be in the form of a Demand Draft from  a nationalised bank issued in favour of HLL Lifecare Limited, Thiruvananthapuram, or in the form of an irrevocable Bank  guarantee of any Nationalised Bank , which should be placed in a separate sealed cover marked “Earnest Money”  and shall be submitted along with the tenders.

19.2	The earnest money shall be in Indian Rupees.

19.3	The earnest money in the form of BG shall be valid for a period of forty five (45) days beyond the validity period of the tender. As validity period of Tender as per Clause 20 of GIT is 120 days, the EMD shall be valid for 165  days from Techno – Commercial Tender opening date.

19.4	Unsuccessful tenderers’ earnest money will be returned to them without any interest, not later than thirty days after conclusion of the resultant contract. Successful tenderer’s earnest money will be returned without any interest, after receipt of performance security from that tenderer.

19.5	Earnest Money is required to protect the purchaser against the risk of the Tenderer’s conduct, which would warrant the forfeiture of the EMD. Earnest money of a tenderer will be forfeited, if the tenderer withdraws or amends its tender or impairs or derogates from the tender in any respect within the period of validity of its tender or if it comes to notice that the information/documents furnished in its tender is incorrect, false, misleading or forged without prejudice to other rights of the purchaser. The successful tenderer’s earnest money will be forfeited without prejudice to other rights of Purchaser if it fails to furnish the required performance security within the specified period.

20.	Tender Validity

20.1	If not mentioned otherwise, the tenders shall remain valid for acceptance for a period of 120 days after the date of tender opening prescribed in the Tender document. Any tender valid for a shorter period shall be treated as unresponsive and rejected.

20.2	In exceptional cases, the tenderers may be requested by the purchaser to extend the validity of their tenders up to a specified period. Such request(s) and responses thereto shall be conveyed by e-mail or by surface mail. The tenderers, who agree to extend the tender validity, are to extend the same without any change or modification of their original tender and they are also to extend the validity period of the EMD accordingly.

20.3	In case the day up to which the tenders are to remain valid falls on/ subsequently declared a holiday or closed day for the purchaser, the tender validity shall automatically be extended up to the next working day.

21.	Signing and Sealing of Tender

  - The tenderers shall submit their tenders as per the instructions contained in GIT Clause 11

21.2	   The original of the tender shall either be typed or written in indelible ink and the same shall be signed by the tenderer or by a person(s) who has been duly authorized to bind the tenderer to the contract. The letter of authorization shall be by a written and notarised power of attorney executed on non judicial stamp paper of Rs.300/-, which shall also be furnished along with the tender.

21.3 	   All the copies of the tender shall be duly signed at the appropriate places as indicated in the Tender documents and all other pages of the tender including printed literature, if any shall be initialled by the same person(s) signing the tender. The tender shall not contain any erasure or overwriting, except as necessary to correct any error made by the tenderer and, if there is any such correction; the same shall be initialled by the person(s) signing the tender.

21.4	   The tenderer is to seal the tender in envelopes, and write the address of the purchaser and the tender reference number on the envelopes. The sentence “NOT TO BE OPENED” before 29.07.2019, 15.00 Hrs (The tenderer is to put the date & time of tender opening) are to be written on these envelopes. The inner envelopes are then to be put in a bigger outer envelope, which will also be duly sealed, marked etc. as above. If the outer envelope is not sealed and marked properly as above, the purchaser will not assume any responsibility for its misplacement, premature opening, late opening etc.

21.5	   Tender document seeks quotation following Two Bid System, in two parts. First part will be known as ‘Techno - Commercial Tender’, and the second part ‘Price Tender’ as specified in clause 11 of GIT. Tenderer shall seal ‘Techno - Commercial Tender’ and ‘Price Tender’ separately and covers will be suitably super scribed. Both these sealed covers shall be put in a bigger cover and sealed.

D. SUBMISSION OF TENDERS

22.	Submission of Tenders

  - Tenders shall be submitted to HLL Lifecare Ltd., by hand or through registered post or courier service at the address given below and not later than 14.00 Hrs on 08.07.2019. In respect of Applications received by post or courier, HLL shall not assume any responsibility for any delayed delivery. Documents submitted in connection with this tender will be treated confidential.

### The Tenders should be addressed to:

Deputy General Manager (Materials),

HLL Lifecare Limited (A Govt. of India Enterprise)

Akkulam  Factory, Sreekariam  P.O.

Thiruvananthapuram – 17

Kerala, India.

22.3   Purchaser may at its discretion, extend the deadline for the submission of Tender, in which case all rights and obligations of Purchaser and the Tenderer subject to the previous deadline shall thereafter be subject to the deadline as extended.

23.	Late Tender

23.1	A tender, which is received after the specified date and time for receipt of tenders will be treated as “late” tender and will be ignored.

24.	Alteration and Withdrawal of Tender

  - The tenderer, after submitting its tender, is permitted to withdraw/alter/modify its tender so long as such withdrawal/alterations /modifications are received duly signed, sealed and marked like the original tender, within the deadline for submission of tenders. Alterations / modifications to tenders received after the prescribed deadline will not be considered.

  - No tender should be withdrawn after the deadline for submission of tender and before

expiry of the tender validity period. If a tenderer withdraws the tender during this period, it will result in forfeiture of the earnest money furnished by the tenderer in its tender.

E.    TENDER OPENING

25. 	Opening of Tenders

25.1   The tenders will be opened at the specified date and time and at the specified place as indicated in the Schedule for submission of tenders.

In case the specified date of tender opening falls on/is subsequently declared a holiday or closed day for the purchaser, the tenders will be opened at the appointed time and place on the next working day.

25.2	Authorized representatives of the tenderers, who have submitted tenders on time may attend the tender opening provided they bring with them letters of authority from the corresponding tenderers.

25.3	Two -Bid Tender system as mentioned in para 21.5 above will be as follows. The Techno - Commercial Tenders are to be opened in the first instance, at the prescribed time and date as indicated in NIT. These Tenders shall be scrutinized and evaluated by the competent committee/ authority with reference to parameters prescribed in the TE document. During the Techno - Commercial Tender evaluation, the tender opening official(s) will assess the salient features of the tenders like brief description of the goods offered, delivery period, Earnest Money Deposit and any other special features of the tenders, as deemed fit by the tender opening evaluation committee. Thereafter, in the second stage, the Price Tenders of only the Techno - Commercially acceptable offers (as decided in the first stage) shall be opened for further scrutiny and evaluation on a date notified after the evaluation of the Techno – Commercial tender. The prices, special discount if any of the goods offered etc., as deemed fit by tender opening official(s) will be read out. The price bid of bidders who do not qualify based on the evaluation of technical bids shall be returned unopened.

F. SCRUTINY AND EVALUATION OF TENDERS

26.	Basic Principle

26.1	Tenders will be evaluated on the basis of the terms & conditions already incorporated in the Tender document, based on which tenders have been received and the terms, conditions etc. mentioned by the tenderers in their tenders. No new condition will be brought in while scrutinizing and evaluating the tenders.  For evaluating the Techno commercial bid, the purchaser may at its discretion call for demonstration/ presentation/ samples etc.

27. Responsiveness

27.1	The Purchaser will examine the Tenders to determine whether they are complete, whether any computational errors have been made, whether required sureties have been furnished, whether the documents have been properly signed stamped and whether the Tenders are generally in order.

27.2	Prior to the detailed evaluation of Price Tenders, the Purchaser will determine the substantial responsiveness of each Tender to the Tender Document. For purposes of these clauses, a substantially responsive Tender is one, which conforms to all the terms and conditions of the Tender Documents without material deviations. Deviations from, or objections or reservations to critical provisions such as those concerning Performance Security (GCC Clause 5),Warranty (GCC Clause 15), EMD (GIT Clause 19), Taxes& Duties (GCC Clause 20),Force Majeure (GCC Clause 26) and Applicable law (GCC Clause 31) will be deemed to be a material deviation. The Purchaser’s determination of a Tender’s responsiveness is to be based on the contents of the tender itself without recourse to extrinsic evidence.

27.3	If a Tender is not substantially responsive, it will be rejected by the Purchaser and cannot subsequently be made responsive by the Tenderer by correction of the nonconformity.

27.4	The tenders will be scrutinized to determine whether they are complete and meet the essential and important requirements, conditions etc. as prescribed in the Tender document. The tenders, which do not meet the basic requirements, are liable to be treated as non – responsive and will be summarily ignored.

27.5	The following are some of the important aspects, for which a tender shall be declared non – responsive and will be summarily ignored;

(i)	Tender form as per Section IX (signed and stamped) not enclosed

(ii)	Tender is unsigned.

(iii)	Tender validity is shorter than the required period.

(iv)	Required EMD (Amount, validity etc.) have not been provided.

(v)	Tenderer has quoted for goods manufactured by other manufacturer’s/ Distributor’s without the required Manufacturer’s/ Distributor’s Authorisation Form/consortium as per Section XIII.

(vi)	Tenderer has not agreed to give the required performance security.

(vii)	Goods offered are not meeting the tender enquiry specification and in demonstration.

(viii)	Tenderer has not agreed to other essential condition(s) specially incorporated in the tender enquiry like terms of payment, liquidated damages clause, warranty clause, dispute resolution mechanism applicable law.

(ix)		Poor/ unsatisfactory past performance.

(x)	Tenderers who stand deregistered/banned/blacklisted by any Govt. Authorities.

(xi)		Tenderer is not eligible as per GIT Clauses 5.1 & 17.1 and Section VIII

(xii)   	Submission of false information in the Tender.

(xiii)   	Erasure or over-writing in the tender is without initialled by the person(s)      signing the tender as per GIT clause 21.

28.	Minor Infirmity/Irregularity/Non-Conformity

28.1	If during the preliminary examination, the purchaser finds any minor infirmity and/or irregularity and/or non-conformity in a tender, the purchaser may waive the same provided it does not constitute any material deviation and financial impact and, also, does not prejudice or affect the ranking order of the tenderers. Wherever necessary, the purchaser will convey its observation on such ‘minor’ issues to the tenderer by registered/speed post/Courier etc. asking the tenderer to respond by a specified date. If the tenderer does not reply by the specified date or gives evasive reply without clarifying the point at issue in clear terms, that tender will be liable to be ignored.

29	Discrepancies in Prices

29.1	If, in the price structure quoted by a tenderer, there is discrepancy between the unit price and the total price (which is obtained by multiplying the unit price by the quantity), the unit price shall prevail and the total price corrected accordingly, unless the purchaser feels that the tenderer has made a mistake in placing the decimal point in the unit price, in which case the total price as quoted shall prevail over the unit price and the unit price corrected accordingly.

29.2	If there is an error in a total price, which has been worked out through addition and/or subtraction of subtotals, the subtotals shall prevail and the total corrected; and

29.3	If there is a discrepancy between the amount expressed in words and figures, the amount in words shall prevail, subject to sub clause 29.1 and 29.2 above.

29.4	If, as per the judgement of the purchaser, there is any such arithmetical discrepancy in a tender, the same will be suitably conveyed to the tenderer by registered/speed post/Courier. If the tenderer does not agree to the observation of the purchaser, the tender is liable to be ignored.

30.    Qualification Criteria

  - Tenders of the tenderers, who do not meet the required Qualification Criteria prescribed in Section VIII, will be treated as non - responsive and will not be considered further.

31.    Tender currency (Indian Rupee)

31.1	The TE document permits the tenderers to quote their prices in Indian Rupees only.

32.	Items Evaluation

- The List of Requirements contains more than one item. The responsive tenders will be evaluated and compared only on package cost as a turnkey tender. However the quotes shall be with item-wise unit cost and total cost with applicable taxes/duties for each and every equipment in the List of Requirements separately.

33.     Comparison of Tenders

The comparison of the responsive tenders shall be carried out on Free Delivery at                consignee site basis. The price ranking will be carried out as under.

Total price = Price with all accessories as per technical specification in the format given in Section – X (A)

The tenderers shall also quote for CAMC price for 5 years as per Section – X (B) separately.  Failing which the bid is liable to be rejected.

34.	Additional Factors and Parameters for Evaluation and Ranking of Responsive Tenders

34.1	Further to GIT Clause 34 above, the purchaser’s evaluation of a tender will include and take into account the following:

In the case of goods manufactured in India or goods of foreign origin already located in India, GST & other similar applicable additional charges etc which will be contractually payable by the tenderer.

  - The purchaser’s evaluation of tender will also take into account the additional factors, if any, incorporated in other Sections, in the manner and to the extent indicated therein.

35.	Tenderer’s capability to perform the contract

35.1	The purchaser, through the above process of tender scrutiny and tender evaluation will determine to its satisfaction whether the tenderer, whose tender has been determined as the lowest evaluated responsive tender is eligible, qualified and capable in all respects to perform the contract satisfactorily.

35.2	The above-mentioned determination will, interalia, take into account the tenderer’s financial, technical and production/execution capabilities for satisfying all the requirements of the purchaser as incorporated in the TE document. Such determination will be based upon scrutiny and examination of all relevant data and details submitted by the tenderer in its tender as well as such other allied information as deemed appropriate by the purchaser.

36.	Contacting the Purchaser

36.1	From the time of submission of tender to the time of awarding the contract, if a tenderer needs to contact the purchaser for any reason relating to this tender enquiry and / or its tender, it should do so only in writing.

36.2	In case a tenderer attempts to influence the purchaser in the purchaser’s decision on scrutiny, comparison & evaluation of tenders and awarding the contract, the tender of the tenderer shall be liable for rejection in addition to appropriate administrative actions being taken against that tenderer, as deemed fit by the purchaser.

G.   AWARD OF CONTRACT

37.	Purchaser’s Right to accept any tender and to reject any or all tenders

  - The purchaser reserves the right to accept in part or in full any tender or reject any or             more tender(s) without assigning any reason or to cancel the tendering process and reject all tenders at any time prior to award of contract, without incurring any liability, whatsoever to the affected tenderer or tenderers.

38.	Award Criteria

- The contract will be awarded to the lowest evaluated responsive tenderer decided by the purchaser as per clause 33 to 36 of GIT.

Only those bidders who qualify at the techno- commercial stage will be eligible for opening of price bids.

39.	Variation of Quantities at the Time of Award/ Currency of Contract

39.1	At the time of awarding the contract, the purchaser reserves the right to increase or decrease the quantity of goods and services mentioned in the schedule (s) in the “List of Requirements” without any change in the unit price and other terms & conditions quoted by the tenderer.

39.2	If the quantity has not been increased at the time of the awarding the contract, the purchaser reserves the right to increase the quantity of goods and services mentioned in the contract without any change in the unit price and other terms & conditions mentioned in the contract, during the currency of the contract period.

40.	Notification of Award

40.1	Before expiry of the tender validity period, the purchaser will notify the successful tenderer(s) in writing, by registered / speed post or by email that its tender for goods & services, which have been selected by the purchaser, has been accepted, also briefly indicating therein the essential details like description, specification and quantity of the goods & services and corresponding prices accepted. The successful tenderer must furnish to the purchaser the required Performance Security within Thirty days from the date of dispatch of the notification of award, failing which the EMD will be forfeited and the award will be cancelled. Relevant details about the performance security have been provided under GCC Clause 5 under Section III

41.	Issue of Contract

41.1	Promptly after notification of award, the Purchaser will mail the Contract Form (as per Section XIV and XV) duly completed and signed, in duplicate, to the successful tenderer by registered / speed post.

41.2	Within twenty one days from the date of the contract, the successful tenderer shall return the original copy of the contract, duly signed and dated, to the Purchaser by registered / speed post.

42.	Non-receipt of Performance Security and Contract by the Purchaser

42.1	Failure of the successful tenderer in providing performance security and / or returning contract copy duly signed in terms of GIT clauses 41 and 42 above shall make the tenderer liable for forfeiture of its EMD and, also, for further actions by the Purchaser against it as per the clause 24 of GCC – Termination of default.

43.	Return of EMD

43.1	The earnest money of the successful tenderer and the unsuccessful tenderers will be returned to them without any interest, whatsoever, in terms of GIT Clause 19.4 and 19.5.

44.	Corrupt or Fraudulent Practices

44.1	It is required by all concerned namely the Consignee/ Tenderers/ Suppliers/ Consortiumetc to observe the highest standard of ethics during the procurement and execution of such contracts. In pursuance of this policy, the Purchaser: -

defines, for the purposes of this provision, the terms set forth below as follows:

“corrupt practice” means the offering, giving, receiving or soliciting of anything of   value to influence the action of a public official in the procurement process or in contract execution; and

“fraudulent practice” means a misrepresentation of facts in order to influence a procurement process or the execution of a contract to the detriment of the Purchaser, and includes collusive practice among Tenderers (prior to or after Tender submission) designed to establish Tender prices at artificial non-competitive levels and to deprive the Purchaser of the benefits of free and open competition;

will reject a proposal for award if it determines that the Tenderer recommended for award has engaged in corrupt or fraudulent practices in competing for the contract in question;

will declare a firm ineligible, either indefinitely or for a stated period of time, to be awarded a contract by the purchaser if it at any time determines that the firm has engaged in corrupt or fraudulent practices in competing for, or in executing the contract.

SECTION - III

GENERAL CONDITIONS OF CONTRACT (GCC)

TABLE OF CLAUSES

| Sl No. | Topic | Page No. |
| --- | --- | --- |
| 1 | Application | 23 |
| 2 | Use of contract documents and information | 23 |
| 3 | Patent Rights | 23 |
| 4 | Country of Origin | 23 |
| 5 | Performance Security | 24 |
| 6 | Technical Specifications and Standards | 24 |
| 7 | Packing and Marking | 24 |
| 8 | Inspection, Testing and Quality Control | 25 |
| 9 | Terms of Delivery | 26 |
| 10 | Transportation of Goods | 26 |
| 11 | Insurance | 26 |
| 12 | Spare parts | 27 |
| 13 | Incidental services | 27 |
| 14 | Distribution of Dispatch Documents for Clearance/Receipt of Goods | 28 |
| 15 | Warranty | 28 |
| 16 | Assignment | 29 |
| 17 | Sub Contracts | 29 |
| 18 | Modification of contract | 29 |
| 19 | Prices | 30 |
| 20 | Taxes and Duties | 30 |
| 21 | Terms and mode of Payment | 30 |
| 22 | Delay in the supplier’s performance | 32 |
| 23 | Liquidated Damages | 33 |
| 24 | Termination for default | 33 |
| 25 | Termination for insolvency | 33 |
| 26 | Force Majeure | 34 |
| 27 | Termination for convenience | 34 |
| 28 | Governing language | 34 |
| 29 | Notices | 35 |
| 30 | Resolution of disputes | 35 |
| 31 | Applicable Law | 35 |
| 32 | General/Miscellaneous Clauses | 35 |

SECTION - III

GENERAL CONDITIONS OF CONTRACT (GCC)

1.	Application

1.1	The General Conditions of Contract incorporated in this section shall be applicable for this purchase to the extent the same are not superseded by the Special Conditions of Contract prescribed under Section IV, List of requirements under Section V and Technical Specification under Section VI of this document.

2.	Use of contract documents and information

2.1	The supplier shall not, without the purchaser’s prior written consent, disclose the contract or any provision thereof including any specification, drawing, sample or any information furnished by or on behalf of the purchaser in connection therewith, to any person other than the person(s) employed by the supplier in the performance of the contract emanating from this TE document. Further, any such disclosure to any such employed person shall be made in confidence and only so far as necessary for the purposes of such performance for this contract.

2.2	Further, the supplier shall not, without the purchaser’s prior written consent, make use of any document or information mentioned in GCC sub-clause 2.1 above except for the sole purpose of performing this contract.

2.3	Except the contract issued to the supplier, each and every other document mentioned in GCC sub-clause 2.1 above shall remain the property of the purchaser and, if advised by the purchaser, all copies of all such documents shall be returned to the purchaser on completion of the supplier’s performance and obligations under this contract.

3.	Patent Rights

3.1	The supplier shall, at all times, indemnify and keep indemnified the purchaser, free of cost, against all claims which may arise in respect of goods & services to be provided by the supplier under the contract for infringement of any intellectual property rights or any other right protected by patent, registration of designs or trademarks. In the event of any such claim in respect of alleged breach of patent, registered designs, trade marks etc. being made against the purchaser, the purchaser shall notify the supplier of the same and the supplier shall, at his own expenses take care of the same for settlement without any liability to the purchaser.

4.	Country of Origin

4.1	All goods and services to be supplied and provided for the contract shall have the origin in India or in the countries with which the Government of India has trade relations.

4.2	The word “origin” incorporated in this clause means the place from where the goods are, manufactured, or from where the services are arranged.

- The country of origin may be specified in the Price Schedule.

5.	Performance Security

5.1	Within thirty (30) days from date of the issue of notification of award by the Purchaser, the supplier, shall furnish Performance Security to the Purchaser for an amount equal to five percent (5%) of the total value of the contract, valid up to 60 days after the date of completion of 3 years warranty period / all contractual obligations by the supplier, including the warranty obligations, initially valid for a total period of minimum 38 months from the date of Notification of Award.

5.2	The Performance security shall be denominated in Indian Rupees.

a)	It shall be in any one of the forms namely Account Payee Demand Draft drawn from any Nationalised bank in India or Bank Guarantee issued by a Nationalised bank in India, in the prescribed form as provided in section XII of this document in favour of the Purchaser. The validity of Bank Guarantee will be for a period up to sixty (60) days beyond respective Warranty Period.

5.3	In the event of any failure /default of the supplier with or without any quantifiable loss to the Purchaser the amount of the performance security is liable to be forfeited. The Purchaser may do the needful to cover any failure/default of the supplier with or without any quantifiable loss to the Purchaser.

5.4	In the event of any amendment issued to the contract, the supplier shall, within twenty-one (21) days of issue of the amendment, furnish the corresponding amendment to the Performance Security (as necessary), rendering the same valid in all respects in terms of the contract, as amended.

5.5	The supplier shall enter into CAMC as per the ‘Contract Form – B’ in Section XV with Purchaser/Purchaser’s Client, 3 (three) months prior to the completion of Warranty Period. The CAMC will commence from the date of expiry of the Warranty Period.

5.6	Subject to GCC sub – clause 5.3 above, the Purchaser will release the Performance Security without any interest to the supplier on completion of the supplier’s  warranty and contractual obligations including submission of satisfactory performance certificates received from Hospital authorities towards quarterly preventive maintenance and breakdown maintenance services rendered by the contractor during warranty  period.

6.	Technical Specifications and Standards

6.1	The Goods & Services to be provided by the supplier under this contract shall conform to the technical specifications and quality control parameters mentioned in ‘Technical Specification’ and ‘Quality Control Requirements’ under Sections VI and VII of this document.

7.	Packing and Marking

7.1	The packing for the goods to be provided by the supplier should be strong and durable enough to withstand, without limitation, the entire journey during transit including transhipment (if any), rough handling, open storage etc. without any damage, deterioration etc. As and if necessary, the size, weights and volumes of the packing cases shall also take into consideration, the remoteness of the final destination of the goods and availability or otherwise of transport and handling facilities at all points during transit up to final destination as per the contract.

7.2	The quality of packing, the manner of marking within & outside the packages and provision of accompanying documentation shall strictly comply with the requirements as provided in Technical Specifications and Quality Control Requirements under Sections VI and VII and in SCC under Section IV. In case the packing requirements are amended due to issue of any amendment to the contract, the same shall also be taken care of by the supplier accordingly.

7.3	Packing instructions:

Unless otherwise mentioned in the Technical Specification and Quality Control Requirements under Sections VI and VII and in SCC under Section IV, the supplier shall make separate packages for each consignee (in case there is more than one consignee mentioned in the contract) and mark each package on three sides with the following with indelible paint of proper quality:

- a. contract number and date

- b. brief description of goods including quantity

- c. packing list reference number

- d. country of origin of goods

- e. consignee’s name and full address and

- f. supplier’s name and address

8.	Inspection, Testing and Quality Control

8.1	The purchaser and/or its nominated representative(s) will inspect and/or test the ordered goods and the related services to confirm their conformity to the contract specifications and other quality control details incorporated in the contract. The purchaser shall inform the supplier in advance, in writing, the purchaser’s programme for such inspection and, also the identity of the officials to be deputed for this purpose. The cost towards the transportation, boarding & lodging will be borne by the purchaser and/or its nominated representative(s).

8.2	The Technical Specification and Quality Control Requirements incorporated in the contract shall specify what inspections and tests are to be carried out and, also, where and how they are to be conducted. If such inspections and tests are conducted in the premises of the supplier all reasonable facilities and assistance, including access to relevant drawings, design details and production data, shall be furnished by the supplier to the purchaser’s inspector at no charge to the purchaser.

- 8.3 	If during such inspections and tests the contracted goods fail to conform to the required specifications and standards, the purchaser’s inspector may reject them and the supplier shall either replace the rejected goods or make all alterations necessary to meet the specifications and standards, as required, free of cost to the purchaser and resubmit the same to the purchaser’s inspector for conducting the inspections and tests again.

8.4	In case the contract stipulates pre-despatch inspection of the ordered goods at supplier’s premises, the supplier shall put up the goods for such inspection to the purchaser’s inspector well ahead of the contractual delivery period, so that the purchaser’s inspector is able to complete the inspection within the contractual delivery period.

8.5	If the supplier tenders the goods to the purchaser’s inspector for inspection at the last moment without providing reasonable time to the inspector for completing the inspection within the contractual delivery period, the inspector may carry out the inspection and complete the formality beyond the contractual delivery period at the risk and expense of the supplier.

8.6	The purchaser’s/consignee’s contractual right to inspect, test and, if necessary, reject the goods after the goods’ arrival at the final destination shall have no bearing of the fact that the goods have previously been inspected and cleared by purchaser’s inspector during pre-despatch inspection mentioned above.

8.7	Goods accepted by the Purchaser and/or its inspector at initial inspection and in final inspection in terms of the contract shall in no way dilute purchaser’s/consignee’s right to reject the same later, if found deficient in terms of the warranty clause of the contract, as incorporated under GCC Clause 15.

- If required by the purchaser, Foreign supplier shall also have the equipment inspected by recognised/ reputed agency like SGS, Lloyd or equivalent (acceptable to the purchaser) prior to despatch at the supplier’s cost and furnish necessary certificate from the said agency in support of their claim.

- For details of final inspection please refer section IV, special conditions of Contract.

9.0 	Terms of Delivery

- Goods shall be delivered by the supplier in accordance with the terms of delivery as follows:

| SI. No. | Decription of work | Completion Period from the date of order |
| --- | --- | --- |
| 1 | Setting up of Microbiology – Phase 1 | 90 days |
| 2 | Providing Consultancy Services for NABL accreditation – Phase 2 | 180 days |

10. 	Transportation of Goods

The supplier shall at their own expenses, arrange transport (including air/sea/land), loading & unloading of goods upto the consignee address.

11.	Insurance:

11.1	Unless otherwise instructed in the SCC, the supplier shall make arrangements for insuring the goods against loss or damage incidental to manufacture or acquisition, transportation, storage and delivery in the following manner:

In case of supply of domestic goods on Consignee site basis, the supplier shall be responsible till the entire goods contracted for arrival in good condition at destination. The transit risk in this respect shall be covered by the Supplier by getting the goods duly insured. The insurance cover shall be obtained by the Supplier and should be valid till installation, testing and commissioning of  the equipments.

If the equipments are not commissioned and handed over to the consignee within stipulated period, the insurance will be  extended by the supplier at their cost till the successful installation, testing, commissioning and handing over of the goods to the consignee. In case the delay in the installation and commissioning is due to handing over of the site to the supplier by the consignee, such extensions of the insurance will still be done by the supplier, but the insurance extension charges at actuals will be reimbursed.

12.	Spare parts

12.1	If specified in the List of Requirements and in the resultant contract, the supplier shall supply/provide any or all of the following materials, information etc. pertaining to spare parts manufactured and/or supplied by the supplier:

a)	The spare parts as selected by the Purchaser to be purchased from the supplier, subject to the condition that such purchase of the spare parts shall not relieve the supplier of any contractual obligation including warranty obligations; and

b)	In case the production of the spare parts is discontinued:

- Sufficient advance notice to the Purchaser before such discontinuation to provide adequate time to the purchaser to purchase the required spare parts, etc. at the supplier’s risk and cost ., and

ii)	Immediately following such discontinuation, providing the Purchaser, free of cost, the designs, drawings, layouts and specifications of the spare parts, as and if requested by the Purchaser.

12.2	Supplier shall carry sufficient inventories to assure ex-stock supply of consumable spares for the goods so that the same are supplied to the Purchaser promptly on receipt of order from the Purchaser.

13.	Incidental services

13.1	Subject to the stipulation, if any, in the SCC (Section – IV), List of equipments (Section – V) and the Technical Specification (Section – VI), the supplier shall be required to perform the following services.

i)	Installation & Commissioning, Supervision and Demonstration of the goods and rectification of accidental damages occurred before handing over the system/site to Hospital authorities.

ii)	Providing required  jigs and tools for assembly, minor civil works required for the completion of the installation.

iii)	Training to end users on equipments in clinical aspects for operating and maintaining the equipments.

iv)	Supplying required number of operation & maintenance manual for the goods.

v)	Providing all the necessary as built drawings after the installation and commissioning.

ix)    	Provide all software updates during warranty period without any additional cost.

14.	Distribution of Dispatch Documents for Clearance/Receipt of Goods

The supplier shall send all the relevant despatch documents well in time to the Purchaser to enable the Purchaser clear or receive (as the case may be) the goods in terms of the contract.

Unless otherwise specified in the SCC, the usual documents involved and the drill to be followed in general for this purpose are as follows.

A) 	For Domestic Goods, including goods already imported by the supplier under its own arrangement.

Within 24 hours of despatch, the supplier shall notify the purchaser, consignee, and others concerned if mentioned in the contract, the complete details of despatch and also supply the following documents to them by registered post / speed post (or as instructed in the contract):

Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount;

Consignee Receipt Certificate as per Section XVI in original issued by the authorized representative of the consignee;

Two copies of packing list identifying contents of each package;

Inspection certificate issued by the nominated Inspection agency, if any.

Certificate of origin;

Insurance Certificate as per GCC Clause 11

Manufacturers/Supplier’s warranty certificate & In-house inspection certificate.

15.	Warranty

15.1	The supplier warrants comprehensively that the goods supplied under the contract is new, unused and incorporate all recent improvements in design and materials unless prescribed otherwise by the purchaser in the contract. The supplier further warrants that the goods supplied under the contract shall have no defect arising from design, materials (except when the design adopted and / or the material used are as per the Purchaser’s/Consignee’s specifications) or workmanship or from any act or omission of the supplier, that may develop under normal use of the supplied goods under the conditions prevailing in India.

- 15.2	This warranty shall remain valid for 3 years as per SECTION V (LIST OF REQUIREMENTS) from the date of handing over entire equipment and acceptance by the Purchaser/Hospital in terms of the contract after installation and commissioning, unless specified otherwise in the SCC.

15.3	In case of any claim arising out of this warranty, the Purchaser shall promptly notify the same in writing to the supplier. The period of the warranty will be as per GCC clause number 15.2 above irrespective of any other period mentioned elsewhere in the bidding documents.

15.4	Upon receipt of such notice, the supplier shall, within 8 hours on a 24 (hrs) X 7 (days) X 365 (days) basis respond to take action to repair or replace the defective goods or parts thereof, free of cost, at the ultimate destination. The supplier shall take over the replaced parts/goods after providing their replacements and no claim, whatsoever shall lie on the purchaser for such replaced parts/goods thereafter. The penalty clause form on rectification will be applicable as per tender conditions.

15.5	In the event of any rectification of a defect or replacement of any defective goods during the warranty period, the warranty for the rectified/replaced goods shall be extended to a further period from the date such rectified / replaced goods starts functioning to the satisfaction of the purchaser.

15.6	If the supplier, having been notified, fails to respond to take action to repair or replace the defect(s) within 8 hours on a 24(hrs) X 7 (days) X 365 (days) basis, the purchaser may proceed to take such remedial action(s) as deemed fit by the purchaser, at the risk and expense of the supplier and without prejudice to other contractual rights and remedies, which the purchaser may have against the supplier.

15.7	During Warranty period, the supplier is required to visit  consignee’s site at least once in 4 months commencing from the date of the installation for preventive maintenance of the goods.

15.8	The Purchaser reserve the rights to enter into CAMC between Purchaser/Purchaser’s Client and the Supplier for the period of 5 years as mentioned in this tender, after the completion of 5 year warranty.

15.9	The supplier along with its Indian Agent and the CAMC provider shall ensure continued supply of the spare parts for the constant performance of the machines and equipments supplied by them to the purchaser for 10 years from the date of installation and handing over.

15.10	The Supplier along with its Indian Agent and the CAMC Provider shall always accord most favoured client status to the Purchaser vis-à-vis its other Clients/Purchasers of its equipment/machines/goods etc. and shall always give the most competitive price for its machines/equipment supplied to the Purchaser.

16.	Assignment

16.1	The Supplier shall not assign, either in whole or in part, its contractual duties, responsibilities and obligations to perform the contract, except with the Purchaser’s prior written permission.

17.	Sub Contracts

17.1	The Supplier shall notify the Purchaser in writing of all sub contracts awarded under the contract if not already specified in its tender. Such notification, in its original tender or later, shall not relieve the Supplier from any of its liability or obligation under the terms and conditions of the contract.

17.2	Sub contract shall be only for bought out items and sub-assemblies.

17.3	Sub contracts shall also comply with the provisions of GCC Clause 4 (“Country of Origin”).

18.	Modification of contract

18.1	If necessary, the purchaser may, by a written order given to the supplier at any time during the currency of the contract, amend the contract by making alterations and modifications within the general scope of contract in any one or more of the following:

a) Specifications, drawings, designs etc. where goods to be supplied under the contract are to be specially manufactured for the purchaser,

b) Mode of packing,

c) Incidental services to be provided by the supplier

d) Mode of despatch,

e) Place of delivery, and

f)  Any other area(s) of the contract, as felt necessary by the purchaser depending on the merits of the case.

18.2	In the event of any such modification/alteration causing increase or decrease in the cost of goods and services to be supplied and provided, or in the time required by the supplier to perform any obligation under the contract, an equitable adjustment shall be made in the contract price and/or contract delivery schedule, as the case may be, and the contract  shall be amended accordingly.

19.	Prices

19.1	Prices to be charged by the supplier for supply of goods and provision of services in terms of the contract shall not vary from the corresponding prices quoted by the supplier in its tender and incorporated in the contract, except for any price adjustment authorised in the SCC.

20.	Taxes and Duties

20.1	Supplier shall be entirely responsible for all taxes, duties, fees, levies etc. incurred until delivery of the contracted goods to the purchaser.

  - Further instruction, if any, shall be as provided in the SCC.

21.	Terms and Mode of Payment

- Phase-1

- Payment Terms  -  Equipment Supply

## a)	50% payment of the Phase – 1 order value shall be released against supply and certificate for receipt of the item in good condition from the end user/purchaser.

## b)   45% payment of the Phase -1 order value shall be released against certificate of installation and commissioning of equipments and turnkey works certified by the end user/purchaser.

c)     Final 5% of the Phase-1 order value shall be released against submission of certificate issued by end user/purchaser certifying that the facility has been installed, commissioned and handed over and submission on Performance Security and other documents stated in this tender.

d)    Payment shall be made subject to recoveries, if any, by way of liquidated damages or any other charges as per terms & conditions of contract and upon the submission of the following documents:

i.  Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount;

ii. Consignee Receipt Certificate as per Section XVII in original issued by the authorized representative of the consignee;

iii. Two copies of packing list identifying contents of each package;

iv.  Inspection certificate issued by the nominated Inspection agency, if any.

v.  Insurance Certificate as per GCC Clause 11 and documents also to be submitted for payment of LC confirming that dispatch documents has already been sent to all concerned as per the contract within 24 hours;

vi.  Certificate of origin.

- Payment Terms  - Turnkey part (Civil & Electrical works)

- 50% of the total value of turnkey work may be released upon the completion of 60 % of the work recommended by State In charge (P&S, HLL) and approved by Project Engineer (P&S, HLL)/ Manager (P&S, HLL).

- Balance 45% will be paid along with the final bill only after submission of Work Completion certificate issued by the concerned hospital and approved by Project Engineer (P&S, HLL)/ Manager (P&S, HLL).

- 5% amount of the Bill will be retained as retention money and will be released only after the Defect liability period of 12 months.

- Final Bill to be submitted detailing the work description, quantity and rate as per the Work Order. Payment will be made against actual measurements recorded and certified jointly by State In Charge (P&S, HLL) and the Contractor ’s representative, counter signed by Project Engineer (P&S, HLL)/ Manager (P&S, HLL). The work completion certificate will be issued by the concerned centres/hospitals and it is the responsibility of the contractor to collect the completion certificate certified by the concerned hospital authority & HLL authorities and submit to HLL for balance 50% payment

- The running account bills are to be submitted detailing the work description, quantity and rate as per the Work Order. Payment will be made against actual measurements recorded and certified jointly by State in charge (P&S, HLL) and the Contractor’s representative. For supply of capital items, duly certified delivery challan/supporting documents such as Warranty Certificates etc. shall be enclosed along with bill.

- Tax Deduction: All statutory deductions like GST, Income Tax, Works Contract Tax, E.S.I., P.F. or any other government-imposed liabilities shall be borne by the Contractor (as applicable at the time of execution of job) and shall be deducted from each bill submitted by the Contractor.

- Phase - 2

- Payment Terms

## a)	100% of the Phase – 2 order value shall be released after final audit and getting NABL Accreditation.

  - CAMC Payment :

The payment of CAMC will be made on half yearly basis, after satisfactory completion of said period on submission of bills with supporting documents, certified by the Hospital authorities, in proof of preventive and breakdown maintenance having rendered in the 4 months. In case the supplier fails carry out the preventive maintenance, in the stipulated period, the purchaser may extend the period of contract to the period of non services or impose the penalty on the supplier at the discretion of the Purchaser. However such non service(s) is/ are by a reason of Hospital, the penalty / extension mentioned above will not be applicable.

21.2	The supplier shall not claim any interest on payments under the contract.

21.3	Where there is a statutory requirement for tax deduction at source, such deduction towards income tax and other tax as applicable will be made from the bills payable to the Supplier at rates as notified from time to time.

21.4	The payment shall be made in Indian Rupees.

21.5	The supplier shall send its claim for payment in writing, when contractually due, along with relevant documents etc., duly signed with date.

21.6	While claiming payment, the supplier is also to certify in the bill that the payment being claimed is strictly in terms of the contract and all the obligations on the part of the supplier for claiming that payment has been fulfilled as required under the contract.

21.7   The payment for the preventive maintenance will be released only on the submission of performance security of 5% for the particular year.

22.	Delay in the supplier’s performance

22.1	The supplier shall deliver of the goods and perform the services under the contract within the time schedule specified by the Purchaser in the List of Requirements and as incorporated in the contract.

22.2	Subject to the provision under GCC clause 26, any unexcused delay by the supplier in maintaining its contractual obligations towards delivery of goods and performance of services shall render the supplier liable to any or all of the following sanctions:

- Imposition of liquidated damages

or

(ii) Forfeiture of its performance security and

(iii) Termination of the contract for default.

22.3	If at any time during the currency of the contract, the supplier encounters conditions hindering timely delivery of the goods and performance of services, the supplier shall promptly inform the Purchaser in writing about the same and its likely duration and make a request to the Purchaser for extension of the delivery schedule accordingly. On receiving the supplier’s communication, the Purchaser shall examine the situation as soon as possible and, at its discretion, may agree to extend the delivery schedule, with or without liquidated damages for completion of supplier’s contractual obligations by issuing an amendment to the contract.

22.4	When the period of delivery is extended due to unexcused delay by the supplier, the amendment letter extending the delivery period shall, interalia contain the following conditions:

(a) 	The Purchaser shall recover from the supplier, under the provisions of the clause 23 of the General Conditions of Contract, liquidated damages on the goods and services, which the Supplier has failed to deliver within the delivery period stipulated in the contract.

(b) 	That no increase in price on account of any ground, whatsoever, including any stipulation in the contract for increase in price on any other ground and, also including statutory increase in or fresh imposition of applicable taxes or on account of any other tax or duty which may be levied in respect of the goods and services specified in the contract, which takes place after the date of delivery stipulated in the contract shall be admissible on such of the said goods and services as are delivered and performed after the date of the delivery stipulated in the contract.

(c) 	But nevertheless, the Purchaser shall be entitled to the benefit of any decrease in price on account of reduction in or remission of applicable taxes or any other duty or tax or levy or on account of any other grounds, which takes place after the expiry of the date of delivery stipulated in the contract.

22.5	The supplier shall not dispatch the goods after expiry of the delivery period. The supplier is required to apply to the Purchaser for extension of delivery period and obtain the same before despatch. In case the supplier dispatches the goods without obtaining an extension, it would be doing so at its own risk and no claim for payment for such supply and / or any other expense related to such supply shall lie against the purchaser.

23.	Liquidated damages

23.1	Subject to GCC clause 26, if the supplier fails to deliver any or all of the goods or fails to perform the services within the time frame(s) incorporated in the contract, the Purchaser  shall, without prejudice to other rights and remedies available to the Purchaser  under the contract, deduct from the contract price, as liquidated damages, a sum equivalent to 0.5% per week of delay or part thereof on delayed supply of goods and/or services until actual delivery or performance subject to a maximum of 10% of the contract price or actual liability of the purchaser due to  delayed supply of goods and/or services/ or total performance of the contractor, whichever is higher. Once the maximum is reached Purchaser may consider termination of the contract as per GCC 24.

During the above-mentioned delayed period of supply and / or performance, the conditions incorporated under GCC sub-clause 22.4 above shall also apply.

24.	Termination for default

24.1	The Purchaser, without prejudice to any other contractual rights and remedies available to it (the Purchaser), may, by written notice of default sent to the supplier, terminate the contract in whole or in part, if the supplier fails to deliver any or all of the goods or fails to perform any other contractual obligation(s) within the time period specified in the contract, or within any extension thereof granted by the purchaser pursuant to GCC sub-clauses 22.3 and 22.4.

24.2	In the event of the Purchaser terminates the contract in whole or in part, pursuant to GCC sub-clause 24.1 above, the Purchaser may procure goods and/or services similar to those cancelled, with such terms and conditions and in such manner as it deems fit and the supplier shall be liable to the Purchaser for the extra expenditure, if any, incurred by the Purchaser for arranging such procurement.

24.3	Unless otherwise instructed by the Purchaser, the supplier shall continue to perform the contract to the extent not terminated.

25.	Termination for insolvency

25.1	If the supplier becomes bankrupt or otherwise insolvent, the purchaser reserves the right to terminate the contract at any time, by serving written notice to the supplier without any compensation, whatsoever, to the supplier, subject to further condition that such termination will not prejudice or affect the rights and remedies which have accrued and / or will accrue thereafter to the Purchaser.

26.	Force Majeure

26.1	Notwithstanding the provisions contained in GCC clauses 22, 23 and 24, the supplier shall not be liable for imposition of any such sanction so long the delay and/or failure of the supplier in fulfilling its obligations under the contract is the result of an event of Force Majeure.

26.2	For purposes of this clause, Force Majeure means an event beyond the control of the supplier and not involving the supplier’s fault or negligence and which is not foreseeable and not brought about at the instance of, the party claiming to be affected by such event and which has caused the non–performance or delay in performance. Such events may include, but are not restricted to, acts of the Purchaser  either in its sovereign or contractual capacity, wars or revolutions, hostility, acts of public enemy, civil commotion, sabotage, fires, floods, explosions, epidemics, quarantine restrictions, strikes excluding by its employees , lockouts excluding  by its management, and freight embargoes.

26.3	If a Force Majeure situation arises, the supplier shall promptly notify the Purchaser in writing of such conditions and the cause thereof within seven days of occurrence of such event. Unless otherwise directed by the Purchaser in writing, the supplier shall continue to perform its obligations under the contract as far as reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the Force Majeure event.

26.4	If the performance in whole or in part or any obligation under this contract is prevented or delayed by any reason of Force Majeure for a period exceeding sixty days, either party may at its option terminate the contract Termination of this agreement for whatever reason shall not affect the obligation/ liabilities of both the parties accrued  hereunder in respect of matters at the time of the agreement.

26.5	In case due to a Force Majeure event the Purchaser is unable to fulfil its contractual commitment and responsibility, the Purchaser will notify the supplier accordingly and subsequent actions shall be taken on similar lines described in above sub-paragraphs.

27.	Termination for convenience

27.1	The Purchaser reserves the right to terminate the contract, in whole or in part for its (Purchaser’s) convenience, by serving 7days written notice on the supplier at any time during the currency of the contract. The notice shall specify that the termination is for the convenience of the Purchaser. The notice shall also indicate interalia, the extent to which the supplier’s performance under the contract is terminated, and the date with effect from which such termination will become effective.

27.2	The goods and services which are complete and ready in terms of the contract for delivery and performance within thirty days after the supplier’s receipt of the notice of termination shall be accepted by the Purchaser following the contract terms, conditions and prices. For the remaining goods and services, the Purchaser may decide:

a)	To get any portion of the balance completed and delivered at the contract terms, conditions and prices; and / or

b)	To cancel the remaining portion of the goods and services and compensate the supplier by paying an agreed amount for the cost incurred by the supplier towards the remaining portion of the goods and services.

28.	Governing language

28.1	The contract shall be written in English language following the provision as contained in GIT clause 4. All correspondence and other documents pertaining to the contract, which the parties exchange, shall also be written accordingly in that language.

29.	Notices

29.1	Notice, if any, relating to the contract given by one party to the other, shall be sent in writing or by cable and confirmed in writing. The procedure will also provide the sender of the notice, the proof of receipt of the notice by the receiver. The addresses of the parties for exchanging such notices will be the addresses as incorporated in the contract.

29.2	The effective date of a notice shall be either the date when delivered to the recipient or the effective date specifically mentioned in the notice, whichever is later.

30.	Resolution of disputes

30.1	If dispute or difference of any kind shall arise between the Purchaser and the supplier in connection with or relating to the contract, the parties shall make every effort to resolve the same amicably by mutual consultations.

If the parties fail to resolve their dispute or difference by such mutual consultations within twenty one days of its occurrence the same shall be referred  for arbitration  to a Sole arbitrator to be appointed by the C&MD,  of the Purchaser Arbitration shall be conducted as per the provisions of the Arbitration and Conciliation Act 1996 or any statutory modification or re-enactment thereof for the time be in force. The arbitrator shall  give a reasoned award. The award of the arbitrator shall be final and binding on the parties to this contract. The venue of arbitration shall be Thiruvananthapuram, Kerala, India and the language of the proceedings shall be English.

The arbitrator shall be requested to give reasoned award.

31.	Applicable Law

The contract shall be governed by and interpreted in accordance with the laws of India for the time being in force.

32.	General/ Miscellaneous Clauses

  - Nothing contained in this Contract shall be constructed as establishing or creating between the parties, i.e. the Supplier/its Indian Agent/CAMC Provider on the one side and the Purchaser on the other side, a relationship of master and servant or principal and agent.

  - Any failure on the part of any Party to exercise right or power under this Contract shall not operate as waiver thereof.

  - The Supplier shall notify the Purchaser of any material change would impact on performance of its obligations under this Contract.

  - Each member/constituent of the Supplier/its Indian Agent/CAMC Provider, in case of consortium shall be jointly and severally liable to and responsible for all obligations towards the Purchaser for performance of contract/services including that of its Associates/Sub Contractors under the Contract.

  - The Supplier/its Indian Agent/CAMC Provider shall at all times, indemnify and keep indemnified the Purchaser against all claims/damages etc. for any infringement of any Intellectual Property Rights (IPR) while providing its services under CAMC or the Contract.

  - The Supplier/its Agent/CAMC Provider shall, at all times, indemnify and keep indemnified the Purchaser against any claims in respect of any damages or compensation payable in consequences of any accident or injury sustained or suffered by its employees or agents or by any other third party resulting from or by any action, omission or operation conducted by or on behalf of the supplier/its associate/affiliate etc.

  - All claims regarding indemnity shall survive the termination or expiry of the contract

SECTION –IV

SPECIAL CONDITIONS OF CONTRACT (SCC)

1.   General requirement for eligibility

- In order to decide the responsiveness of tender, the Purchaser may ask to the tenderer for Demonstration of equipment/system, presentations and sample and the tenderer shall arrange Purchaser’s requirement as and when so asked, failing which the tender shall be deemed as non-responsive.

- The Licences, Certifications, if any, required from the regulatory authorities in India with respect to this tender shall be produced along with the tender.

- All technical details, catalogue, application details, shall be provided along with the tender.

- Signed copy of Tender Document (all pages of Bid documents to be signed & stamped) by the Bidder as token of acceptance of the Terms & Conditions.

- Duly filled, signed and sealed forms as per the Annexure of the tender document.

- Power of attorney in Rs 100/-  in non-judicial stamp paper duly notarized by the authorized signatory to sign  and submit the bid documents

- Sales Tax Clearance certificate(self attested copy)

- Copy of PAN Card.(self attested Copy)

- GST Registration Certificate

- Certificate of incorporation / Memorandum of Article (self attested copy)

- Last 3 financial years audited Profit & Loss, Balance Sheet duly certified by Chartered Accountant.

- EMD as per tender document.

- Tender Fee as per tender document. Undertaking for replacement of complaint / defective items as Annexure …

- One to one compliance statement to technical specification requirements against each item shall be provided along with the tender, with pamphlets/Catalogs.

- Acceptance test should be done at designated hospitals, prior to handing over of equipment.

- All details of pre installation and installation works along with schedules & drawings should be supplied within a week of award of order.

2.  Final Inspection

The final inspection of the Goods will be done by the Technical Committee of  the Purchaser and Hospital Authorities after installation and commissioning of the goods.

3. Warranty:

- 3 years as per SECTION V (LIST OF REQUIREMENTS)  for comprehensive warranty as per Conditions of Contract of the TE document for complete equipment from the date of satisfactory installation, commissioning, trial run & handing over of equipment to Hospital/Institution.

- 95% uptime Warranty of complete equipment with extension of Warranty period by double the downtime period on 24 (hrs) X 7 (days) X 365 (days) basis.

- All software updates should be provided by the supplier free of cost during Warranty period.

4. After Sales Service:

After sales service centre should be available at on 24 (hrs) X 7 (days) X 365 (days) basis. Breakdowns/Complaints should be attended within 8 hrs. An undertaking by the Manufacturer shall accompany the tender that the spares for the equipment shall be available for at least 10 years from the date of supply.

5.  CAMC :

a)  	Tenderer shall offer price for

CAMC rates  separately  in SECTION – X, PRICE SCHEDULE (B)

CAMC shall be awarded at the discretion of Purchaser/Hospital at the end of the warranty period.

b) 	The CAMC shall be for 5 years after completion of prescribed warranty period.

c) 	Preventive Maintenance services during CAMC shall be rendered on quarterly basis with minimum gap between two services shall be not less than 90 days and not more than 115 days. In addition, all breakdown calls shall be attended  to  immediately  and  all  major repairs shall be rectified within 7 calendar days from the date of intimation,  as per tender, failing which, a penalty of Rs.50,000/- per day thereof is leviable until the equipment is repaired and commissioned to the    satisfaction  of   the  end  user,  For  this  purpose  supplier shall  carry   sufficient   inventories to assure prompt replacement of defective parts as per tender

- Breakdown calls shall be attended immediately and major complaints shall be rectified within 7 calendar days from the date of intimation. The breakdown calls shall not be combined with preventive maintenance calls.

- In case the performance of CAMC services is not satisfactory and found   below the 95% uptime level, the Purchaser / Hospital has the right to source the maintenance services from other means/agency at the risk and cost of supplier including termination of contract and legal/penal actions.

- On receipt of CAMC order, the supplier shall furnish performance security for an amount equal to 5% of the CAMC value per annum in the form of Demand Draft or Performance Bank Guarantee, which will be renewed in term with value of every year till completion of CAMC period.

- The cost of CAMC includes preventive maintenance with required testing, calibration as per technical/service/operational manual, labour and spares. The supplier shall undertake preventive maintenance as recommended in the manufacturer’s technical/ service /operational manual, but minimum once in three months during the 5 years CAMC period for preventive maintenance.

- The cost of CAMC may be quoted along with taxes and duties applicable on the date of Tender Opening. The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation, the price will be taken inclusive of such taxes and no claim for the same will be entertained later.

- The payment of CAMC will be made once in every four months after satisfactory completion of said period, duly certified by Hospital authorities, but subject to valid Performance Security.

- There will be 95% uptime warranty during CAMC period on 24 (hrs) X 7 (days) X 365 (days) basis, with penalty, to extend CAMC period by double the downtime period.

- During the CAMC period, all software updates should be provided free of cost.

Section V

LIST OF REQUIREMENTS

List of Equipments

| Sl. No | Product | Qty | Unit |
| --- | --- | --- | --- |
| 1 | AHU | 1 | Job |
| 2 | Anaerobic Gas .Jar | 1 | No. |
| 3 | Antibiotic zone scale | 3 | Nos. |
| 4 | Autoclave Horizontal (79 L) | 1 | No. |
| 5 | Autoclave Vertical | 2 | No. |
| 6 | Automated microbial detection system | 1 | Nos. |
| 7 | Beaker 10 ml | 10 | Nos. |
| 8 | Beaker 50 ml | 10 | Nos. |
| 9 | Beaker 250 ml | 10 | Nos. |
| 10 | Beaker 500 ml | 10 | Nos. |
| 11 | Beaker 1000 ml | 10 | Nos. |
| 12 | Bio safety cabinet Class III-CE certified | 2 | Nos. |
| 13 | BOD Incubator 500 | 1 | No. |
| 14 | BOD Incubator 600 | 1 | No. |
| 15 | BOD Incubator 900 | 1 | No. |
| 16 | Centrifuge Refrigerated | 1 | No. |
| 17 | Electronic Balance (Micro/Semi Micro) | 2 | Nos. |
| 18 | Computer Unit | 2 | Nos. |
| 19 | Conical Flask 250ml | 10 | Nos. |
| 20 | Conical Flask 500ml | 10 | Nos. |
| 21 | Conical Flask 1000ml | 10 | Nos. |
| 22 | Cross Over Bench (Stainless Steel) | 1 | No. |
| 23 | Digital Colony Counter | 2 | Nos. |
| 24 | Dropping Bottle tbr Stains (125 ml) | 5 | Nos. |
| 25 | Electric Loop sterilizer | 1 | No. |
| 26 | Flask flat Bottom (1000 ml) | 10 | Nos. |
| 27 | Fumi gation apparatus | 1 | No. |
| 28 | Garment Shelf (Stainless Steel) | 1 | No. |
| 29 | Heating Block | 2 | Nos. |
| 30 | Hot Air Oven (Digitally Control) | 3 | Nos. |
| 31 | Hygrometer | 4 | Nos. |
| 32 | Incubators | 4 | Nos. |
| 33 | Induction Cook Top | 1 | No. |
| 34 | Lanrinar Air flow cabinet | 2 | Nos. |
| 35 | Measuring Cylinder (10 ml) | 10 | Nos. |
| 36 | Measuring Cylinder (25 ml) | 10 | Nos. |
| 37 | Measuring Cylinder (50 ml) | 10 | Nos. |
| 38 | Measuring Cylinder (100 ml) | 10 | Nos. |
| 39 | Measuring Cylinder (250 ml) | 10 | Nos. |
| 40 | Measuring Cylinder (500 ml) | 10 | Nos. |
| 41 | Measuring Cylinder (1000 ml) | 10 | Nos. |
| 42 | Membrane Filtration Unit With all Accessories | 10 | Nos. |
| 43 | Micrometer Stage | 2 | Nos. |
| 44 | Microscope | 1 | No. |
| 45 | Micrometer Eye Piece | 2 | Nos. |
| 46 | Microwave Oven | 1 | No. |
| 47 | Multichannel Pipette-variable volume 5- 50µl, 20-200µl | 5 | set |
| 48 | Magnetic Stirrer with Bead | 2 | Nos. |
| 49 | Petridish with lid 50x17mm | 20 | Nos. |
| 50 | Petridish with lid 80x17mm | 20 | Nos. |
| 51 | Petridish with lid 100x17 mm | 20 | Nos. |
| 52 | Petridish with lid 50x12 mm | 20 | Nos. |
| 53 | Petridish with lid 80x 15mm | 20 | Nos. |
| 54 | Petridish with lid 90x 15mm | 20 | Nos. |
| 55 | Petridish with lid 100x 15mm | 20 | Nos. |
| 56 | Petridish with lid 150x 25mm | 20 | Nos. |
| 57 | Petridish with lid 200x 30mm | 20 | Nos. |
| 58 | pH Detern'rination apparatus | 3 | Nos. |
| 59 | Volumetric Pipette 1 ml | 20 | Nos. |
| 60 | Volumetric Pipette 2 ml | 20 | Nos. |
| 61 | Volumetric Pipette 5 ml | 20 | Nos. |
| 62 | Volumetric Pipette 10 ml | 20 | Nos. |
| 63 | Volumetric Pipette 20 ml | 20 | Nos. |
| 64 | Volumetric Pipette 25 ml | 20 | Nos. |
| 65 | Pipette Graduated 1 ml | 20 | Nos. |
| 66 | Pipette Graduated 2 ml | 20 | Nos. |
| 67 | Pipette Graduated 5 ml | 20 | Nos. |
| 68 | Pipette Graduated 10 ml | 20 | Nos. |
| 69 | Pipette Graduated 25 ml | 20 | Nos. |
| 70 | Pipette Sterilization Can | 5 | Nos. |
| 71 | Reagent Bottles/ Stopper - 100 cc | 60 | Nos. |
| 72 | Reagent Bottles/ Stopper - 1000 cc | 36 | Nos. |
| 73 | Reagent Bottles/ Stopper - 250 cc | 24 | Nos. |
| 74 | Reagent Bottles/ Stopper - 50 cc | 60 | Nos. |
| 75 | Refrigerator (300 ltr or More) | 3 | Nos. |
| 76 | Single channel Pipette-variable volume 0.5-l0µl, 5-50µl, 50-200µl,200-1000 µl | 5 | set |
| 77 | Slides | 5 | Box |
| 78 | Staining (troug) rod (10 slides) | 60 | Nos. |
| 79 | Stainless Steel Bead (1000pcs per box) | 1 | Box |
| 80 | Table (Stainless Steel) | 3 | Nos. |
| 81 | Test tube l00 x l2 mm | 20 | Nos. |
| 82 | Test tube l50 x l8 mnr | 10 | Nos. |
| 83 | Test tube 75 x 12 mm | 20 | Nos. |
| 84 | Thermometer | 5 | Nos. |
| 85 | Stool (Stainless Steel) | 5 | Nos. |
| 86 | Trinocular Fluorescent microscope with digital colour CD Camera LED | 1 | No. |
| 87 | Ultrapure Water Purifier | 1 | No. |
| 88 | UPS (3 kV) | 2 | Nos. |
| 89 | UV Larnp | 2 | Nos. |
| 90 | Voluntetric Flask 50 ml | 20 | Nos. |
| 91 | Voluntetric Flask 100 ml | 20 | Nos. |
| 92 | Voluntetric Flask 250 ml | 20 | Nos. |
| 93 | Voluntetric Flask 500 ml | 20 | Nos. |
| 94 | Voluntetric Flask 1000 ml | 20 | Nos. |
| 95 | Vortex Mixer | 1 | No. |
| 96 | Water bath | 2 | Nos. |
| 97 | Water bath serological 37 deg | 1 | No. |
| 98 | Water bath serological 56 deg | 1 | No. |
| 99 | Micropipette Tips (variable Volume) 0-200 µl | 20 | Box |
| 100 | Micropipette Tips (variable Volume) 200-1000 µl | 20 | Box |
| 101 | Micropipette Tips (variable Volume) 0-10 µl | 20 | Box |
| 102 | Ultra Sonic bath | 2 | Nos. |
| 103 | Vial Decapping machine | 2 | Nos. |

SECTION VI

Technical Specification

Electronic Weighing balance

- Weighing capacity : 220g

- Readability : 0.001g or 0.1mg

- Standard deviation : 0.001g

- Pan Diameter (mm) : 100 x 100 (with guard)

- Weight in kg: approx. 2.2 kg

- Power supply: 100 – 250VAC, 47-63Hz

Hot Air Oven/Sterilizer Capacity : 90L with 2 shelves Specifications:

- Temperature range: +5°C above Ambient to 250°C

- Temperature Accuracy : ± 1°C

- Mild Steel exterior body with Powder coating

- AISI SS 304 grade stainless steel interior body

- Microprocessor based digital display Temperature controller.

- Provided with air circulating fan for temperature uniformity

- Provided with heating elements on three sides for uniform temperature

- Provided with SS 304 stainless steel adjustable shelves

- Provided with air ventilations

- Provision for monitoring the temperature with a thermometer

- Audio Visual alarm for any temperature deviations

- Power : 220-240V/50-60Hz

- Should have ISO & CE standards

Laminar Airflow

- Cabinet/table Size (WDH): 1200 x 60 x 60 cms (4 x 2 x 2 feet)

- Body: Mild Steel with epoxy powder coating

- Working Surface: AISI SS 304 grade stainless steel

- Filter & pre-filters: Minipleat HEPA & Fine filters

- Filtration Size: Down to 0.3 microns

- Lights: Fluorescent and UV

- Petcock for gas/LPG

- Provided with Magnehelic Gauge for monitoring air pressure

- Extra electrical point, 15amp, Multi-pin Socket

- Fully closing, clear Polycarbonate Door with sash handle adjustable to any height; counterbalanced

- Transparent scratch resistant side panels sheet of 6mm thickness

- Castor wheels for easy movement

- Power requirement : 220-240V, 50-60Hz

Distilled Water Plant–

- Three Stage Pre-filtration System with Booster Pump

- Lab NE 15VF is a direct Tap to Ultrapure system combines production of RO and Ultra Pure water from a single unit with, 15 l/h capacity.

- It Comprises of Pretreatment Module, High Pressure Pump, RO Membrane, Dual Wavelength UV Lamp, EDI Module, Ultra Purification cartridge, Storage tank 30L (Option of 60, 100 L Tank is available), Inbuilt Ultra filtration Cartridge, Point of Use filter, Vent Filter for tank and a Remote Water dispenser.

| Source Water | Potable Tap water |
| --- | --- |
| Conductivity, µS/cm | < 2000 |
| Temperature, °C | 5 to 35 |
| Pressure, psi (bar) | 0.05 – 0.5 Mpa 7 – 72 |
| Hardness | < 300 ppm as CaCO3 |
| Feed Water Chlorine, PPM | < 0.5 |

RO Water (Type III):

| Ionic Rejection | >99% |
| --- | --- |
| Bacteria Rejection | >99% |
| Flow Rate | 15 L/h |

Pure Water (Type II):

| Resistivity at 25°C | 10 MΩ.cm |
| --- | --- |
| TOC | < 30 PPB |
| Flow Rate | 15 L/h |

Ultrapure Water (Type I):

| Resistivity at 25°C | 18.2MΩ.cm |
| --- | --- |
| Conductivity AT 25°C | 0.055us/cm |
| TOC | 1 - 5ppb |
| Endotoxin (Pyrogens) | <0.001EU/ml |
| Particulate(>0.2um) | <1/ml |
| Bacteria | <0.1cfu/ml |
| Rnase Dnase <4 pg/µl | <0.01 ng/ml |
| Manual dispense flow rate | Upto 2.0L/min |
| Automatic dispense volume | 100 – 60000 ml |

Electrical Requirements: Electrical Voltage: 220V+10% Electrical Frequency: 50 HZ Packing Information:

Net Weight: 24kg Water tank (30L): 7kg

External Dimensions (W x D x H): 315 x 525 x 570mm Water tank (30L): 380 x 380 x 595mm

Shipping Weight : 37kg Water tank (30L): 15kg

Shipping Dimensions (WxDxH): 525 x 610 x 770mm Water tank (30L): 520 x 440 x 615

Features: Space-saving design permits benchtop installation or wall mounting with integrated

mounting bracket Easy to operate, Recirculation

Serological Water bath

- Volume : 14Ltr (30 x 25 x 17.5( cm)(WDH)

- Temperature range: +5°C above ambient to 100°C

- Temperature Accuracy: ±1°C @37°C

- Temperature Sensitivity : 0.1°C

- Outer Body: Mild Steel with Epoxy Powder coating

- Inner Body: Stainless Steel of grade AISI SS 304

- Temperature control: Microprocessor based digital display Temperature controller

- Insulation: Glass wool

- Easy to clean Stainless Steel tank

- AISI SS 304 pyramid Lid

- Power : 230V/50 Hz

- Should have ISO and CE certificate

pH Meter with digital display

- Measures pH & mV

- Range: pH - 0 to 14 ; mV 0 to ±1999

- Highly stable & durable

- Temperature compensation facility

- Digital display for pH readings

- Auto polarity and decimal indication

- Accuracy : ±0.01 pH & ±1 mV

- Power : 230V ±10 AC, 50Hz

- Standard unit includes pH meter, pH electrode, Dust cover, & Stand.

Incubator

- Capacity : 90L with 2 shelves

- Temperature Range: 5 ºC above Ambient to 70ºC

- Accuracy: ±0.5 ºC

- Microprocessor based digital temperature controller with PT 100 temperature sensors.

- Digital display for temperature values.

- Insulation: PUF foam/Glass wool

- Interior wall made of AISI SS 304 sheet

- Exterior is Epoxy powder coated Mild Steel make

- Double walled unit with Solid Door

- 2 numbers of Adjustable Perforated AISI SS 304 shelves

- Provided with Memmert type heating with circulating fan for uniform temperature

- Power : 220-240V/50-60Hz (5 amps)

- Meets ISO and CE standards

Programmable control Refrigerated Centrifuge

Hi-speed Refrigerated Micro centrifuge, bench top mode provided with microprocessor control & 99 program memory. The parameter continues read out RCF&RPM value, CFC free cooling system, drive default, LCD display locking system by electronic imbalance indicator & safety with brushless induction motor drive with S.S 304 chamber (without rotot head) 2KVA stabilizer recommended

BOD (Refrigerated) Incubator Temperature Range: 5 ºC to 60 ºC

- Accuracy: ±1ºC

- Microprocessor based digital temperature controller

- Insulation: PUF foam

- Interior wall made of AISI SS 304 sheet

- Exterior is Epoxy powder coated Mild Steel

- CFC-free, hermetically sealed Compressor

- Keyed door lock with magnetic door gasket

- Audio visual alarm when the temperature deviated from the preset    temperature

- Double walled unit with Solid Door

- Light provided inside

- Adjustable Perforated AISI SS 304 shelves

- Provided with Castor Wheels

- Audio Visual alarm for any temperature deviations

- Provided with heating element for higher temperature

- Power : 220-240V/50-60Hz (5 amps)

- Should have ISO and CE certificate

Horizontal Cylindrical High Pressure Sterilizer - Manual Operation

- Chamber Size: 400 dia x 600mm depth

- Chamber Volume: 79 Litres

- Operating Temperature: 121oC

- Operating pressure : 1.2 kg/cm sq

- The sterilizer chamber, back plate and the end ring would be formed out of S.S 316 and the jacket of SS 304.

- The jacket is insulated with glass wool and covered with S.S 304.

- Provided with pressure locking type safety hinged door with silicon gasket.

- The door plate is made of SS 316.

- Shooting bolts and steam generator is made of S.S 304

- Stand is provided which is made of duly powder coated steel

- The unit will be provided with multi port operating valve consisting of 5 positions – off/slow exhaust, vacuum dry, sterilization and fast exhaust

- The unit will also be provided with all safety features like safety valve, vacuum breaker, self sterilizing vacuum dryer, pressure switch, gauges, contactor, water level sensor etc.

- Sterilizer specs as above, is suitable for electrical operation on 400/440 volts, 3 phase, AC supply. Electrical load: 6KW.

- Manufactured as per IS: No. 3829 (Part – I) and should bear certification Mark, ISO 9001:2008,

- ISO 13485:2003 and CE Certified.

Computer

- Processor - dual core or i3 Intel processor or equivalent AMD

- RAM – 4 GB

- Hard Drive - 256 GB or larger solid state hard drive

- Wired/Wireless – Key board and Mouse

- Monitor - LCD with Display Port/HDMI or DVI support

- Operating System - Windows 7 or 10 Home or Professional editions.

Digital Thermometer

- Temperature Range : -50°C to +80°C Accuracy range : +2.5°C

- (Equipment supplied with Calibration certificate with NABL Traceability)

- Automated microbial detection system from

Automated Microbial Detection System

- Automated Continuous Monitoring Blood Culture System with minimum 35-40 sample vial capacity.

- System should be a modular one which can be upgraded as per requirement of the Lab with minimum additional cost.

- True walk away system with a simple user interface and touch screen operations.

- System has LIS communication capability for quick result information availability.

- System is based on sensitive fluorescence/Colorimetric technology for interpretation of results.

- The culture media has strong resin based Antibiotic Removal devices to minimize chances of false negatives due to high antibiotics in specimens and have minimal time to detection of organisms.

- System has Auto Quality Control and Calibration facility to avoid any manual daily maintenance. User intervention for routine QC/calibration should not be required.

- System has special media for Pediatric samples and low volume sterile body fluid samples.

- System is supplied along with on line UPS with 30 minutes back-up.

Automated & identification and Antibiotic sensitivity testing system

- It should be Fully Automated, Walk-away, High-throughput & random-access Bacterial Identification & Antibiotic Susceptibility system.

- System should be able to perform incubation and result-interpretation simultaneously.

- System should have minimum 50 sample positions and it can be upgraded on site as and when required with additional modules.

- System should be able to simultaneously Identification and Drug Susceptibility testing (both ID & AST) for minimum 50 samples at a time.

- Minimum detection time for Identification should be 2- 6 hrs.

- System should have advance Software that maintains consistency of Identification and Susceptibility test results.

- Sensitivity detection should be based on turbidity and redox reaction.

- System should employ Combination of chromogenic & florigenic substrates together with 5 independent time specific databases for highly accurate & faster Identification

- System should employ Redox Reaction for Bacterial Metabolism & Turbidity for Bacterial cell division to determine microbial growth during AST, to reduce major errors

- Test panels should be sealed & should never move inside the system to avoid dislodging of panels from the its position, jam, crack  or leak

- System should be able to estimate true Doubling Dilution and there shouldn’t be any skipped dilution of various antibiotics to estimate true MIC for analysis of delayed resistance.

- MIC interpretation should not be based on computer –simulation technology.

- Estimation of Susceptibility should be based on Doubling Dilution MIC inline with the CLSI guideline.

- System should have test panels with various combinations- i.e. Only Identification Panel, Only Sensitivity Panel, and Identification & Sensitivity Combination Panel.

- Primary panels should be able to identify various resistant mechanisms those may be present in the sample- ESBL, MRSA, VRE, HLAR, b-lactamase-producing Staphylococcus (Penicillinase), Macrolide resistance in Streptococci (Efflux/MLSb), High Level Penicillin resistance in S. pneumoniae, Low Level Penicillin resistance in S. pneumoniae.

- Company should provide a brand new unit with certificate of analysis with compatible UPS system with minimum 30 mins back-up.

Channel variable volume micropipette

- 5-50 µl

- 20-200 µl’

Single Channel Pipette volume

- 0.5-10 µl

- 5-50 µl

- 50-200 µl

- 200-1000 µl

BOD Incubator

- Temperature Range: 5 ºC to 60 ºC

- Accuracy: ±1ºC

- Microprocessor based digital temperature controller

- Insulation: PUF foam

- Interior wall made of AISI SS 304 sheet

- Exterior is Epoxy powder coated Mild Steel

- CFC-free, hermetically sealed Compressor

- Keyed door lock with magnetic door gasket

- Audio visual alarm when the temperature deviated from the preset temperature

- Double walled unit with Solid Door

- Light provided inside

- Adjustable Perforated AISI SS 304 shelves

- Provided with Castor Wheels

- Audio Visual alarm for any temperature deviations

- Provided with heating element for higher temperature

- Power : 220-240V/50-60Hz (5 amps)

- Should have  ISO and CE certified

Biohazard Safety Cabinet Class III with 100% exhaust

- Worktable Size : 4 x 2 x 2 feet (1220 x 610 x 610 mm):

- Provided with 2.5ft stand

- Nominal Air inflow velocity of 105fpm (0.5 m/sec)

- Nominal Down flow velocity of 55fpm (0.3 m/sec)

- AISI SS 304 interior and removable, seamless, dished work surface with lift out knobs

- Epoxy Powder coated Mild Steel sheet outer body

- Intrinsically safe negative pressure

- Control System : Microprocessor based controller with digital display

- Display : LCD Display; exhaust filter and down flow filter pressure measurements, filter and UV lamp time, inflow and down flow velocity, filter life, humidity and temperature,   system working fine etc.

- Front Toughened Sash with Glove box system

- Glove Box system is single piece leak-tested Hypalon glove assemblies which guarantee maximum protection.

- An integrated pass-through with interlocking doors permits materials transfer without risk of environmental contamination.

- An angled cabinet front ensures an ergonomic working posture.

- Cabinet operates at negative pressure relative to the laboratory in order to prevent migration of pathogenic materials out of the work area

- Supply and exhaust 99.99% efficient Minipleat HEPA filters with filtration size down to 0.3 microns.

- Curved stainless steel inlet Grill with Secondary Airflow slots

- Glare-free fluorescent and UV light

- Water and Vacuum/Gas petcocks

- Power: 230V, 50Hz

- Requirement 16 amps three pin plug-in

- Exhaust Outlet included

- Virus Burn Out Included

- Should have  ISO and CE certified

Certifications:

- Filter Integrity

- Air Velocity

Trinocular Fluorescent Microscope

Features : UIS2 Infinity Plan Optics for excellent image flatness, Seidentoph binocular head for IPD adjustment, Anti-fungus treatment for tropical durability, Highpoint paired eyepieces (F.N.20), Abbe condenser with high performance aspheric lenses for bright & uniform illumination throughout the field of view, Window in arm for convenient carrying & Ergonomic design for user convenience, SMPS power supply for flicker free illumination , Rack less specimen stage for user safety and comfort, specimen stage focus-lock prevents any accidental damage to objectives

- LED FLUORESCENCE ATTACHMENTS Tr TRINOCULAR MICROSCOPES

Features: Epi-LED - Reflected LED Fluorescence attachment, free from any alignments and ready to attach.

The use of a high power Epi LED enables increased performance replacing conventional mercury arc-lamps. Light Source Lifetime : Above 30000 hrs., allowing long years of operation without any cost of replaceable lamps.

- MU 2.3 MP Camera (without C – Mount Adapter) for Attaching  Trinocular Microscope

Specifications: 2.3 Mega Pixel CMOS Camera with Sony Exmoor sensor having sensor size of 1/1.9", Resolution: 1920 X 1200, suitable for Fluorescence & Bright Field Imaging, pixel size 3.75 X 3.75 um, Binning 1X1, 2X2 with USB 3 port for faster data transfer, comes with a software for point to Point measurement, auto white balance, auto exposure option, basic setting: Gain, Gamma, Contrast, Auto white balance, optical Port : C-Mount compatible with Windows 7 / 8 (32/ 64 bit).

Packing includes: 2.3 MP Camera, USB 3 cable, calibration slide

AIR HANDLING UNIT

Standard      : As per ISO 14644-1/ HTM/ ASHRAE standards

Class Level: ISO 7, Temperature: 23 Deg C ( +/- 2) RH Level: 55 % (+/- 5%)

| SI. No | Description |
| --- | --- |
| 1. | CRFS Factory built package type classified environment system, suitable for aseptic application with built arrangement to produce thermo hygrometric conditions (Temperature, Humidity etc). Also it is suitable to control V.O.C positive pressure. This system is DX type with of cooling capacity and free air delivery of 2000 CFM with 150mm of water column air static. Unit made of panels consisting of hot dip G.I casing of thickness of 24 swg thick for both outside and inside layers encasing polyurethane foam insulation having a density of 38-40 Kg/cum factory injected by using injection molding machine. Unit equipped with EU 4, viscous washable synthetic type, bullet expanded inner grooved coil section of Direct Expansion, 6 Row deep for maximum dehumidification, and Aluminium fins with Hydraufillic coating, coil size shall be selected for a maximum face velocity of 500 feet/minute. Coils shall have automatic air vents, the vent outlets beings piped to the drain pan with a copper pipe. drain cock with valve with aluminium finned blower sections with DIDW blower with motor and drive assembly, belt driven package (V-Belt) with TEFC motor suitable for operations on 415 + 10%, 50 Hz, A.C supply with vibration isolator, fine filter section with EU 7filter with 300mm thick, dampers to be of extruded aluminium low leakage aero file design. VFD for Continuous operation will be provided. |
| 2. | PPGI RIGID WALL PANELING & FALSE CEILING Supply fabricating & fixing of PPGI PUF filled panels. There will be separate supporting structure to which the prefabricated tray formed panels will be assembled. The PPGI panels will have one side pre coated galvanized sheet with PUF sandwiched 45 kg/m3. The corners will size reinforced Composite Panel from behind. The vertical joints will be cold welded and made seamless and aero-sprayed. The PPGI surface will be free from any projection or cavity. |
| 3. | Ducting & Insulation Supply, installation, testing and balancing of Air Distribution System consists of factory fabricated Galvanized Iron sheet metal ducting with TDF/TDC flanges complete with supports (anchor fasteners and tie-rods, cradles, angle iron supports, GI bolts and nuts, 4mm thick self-adhesive neoprene rubber gasket, stiffeners, zinc chromate painted angle iron flanges, chipping and plastering of openings made for duct supports etc., filling the gap with fibre glass wool and with sealant complete as per specification. The rates quoted shall include turning vanes, straightening vanes, splitters, inspection doors probe, inlets with test openings enclosures complete with locking caps, chain, gaskets, insulating plugs, and extension for insulated ductwork etc. Supply and erection of Thermal Insulation - Material shall be FM approved, open cell Nitrile Rubber foam, temperature range from -10ºC to +60ºC, Class O fire category and as per specification. The Thickness of the insulation material shall be 13/19 mm. |
| 4. | CONDENSING UNIT: Supply of 8.5 TR capacity air cooled scroll condensing unit with Hermetic compressors and eco-friendly Refrigerant. All the copper line distribution with Drier, Shut off valves, Expansion valve, etc. |
| 5. | HEPA Filter with stainless steel module : The casing will be made out of Aluminium with Media made out of Micro fibre glass paper and the bonding will be of PU resin. Fins will be Aluminium fins – blended type. Poly ethylene Gasket will be provided. Pressure Drop will be Within 25mm WC. Performance Efficiency will be 99.997% and Grade will be H 13 / H 14 & Penetration - 0.3 sub micron. Terminal module made Stainless steel/Aluminum finish, with bottom grill out of Stainless steel grill with Glossy finish, H 13 HEPA Filter with precise bolt arrangement, will be encased in casing and fixed rigidly into the modular without airleaks, also accessories such as air entry flange connection and D.O.P charging port. Designed for uniform laminar flow across the surface |
| 6. | Electrical Control Panel: Electrical control panel consist of various sets of controllers and VFD. VFD is a general purpose drive which controls AC motors up to 22 KW with maximum reliability complies with RoHS directives. The drive is easy to install PLC controlled and with automatic motor adaptation technology with high efficiency. |
| 7. | Testing & Validation DQ, IQ, OQ & PQ Air velocity test Air changes test Air qty. check test Temp. & RH test DOP test for Hepa filter Particle count test Air flow pattern test Recovery test |

TURNKEY WORK SCHEDULE

| S.No | DESCRIPTION | QTY | UNIT |
| --- | --- | --- | --- |
|  | DISMANTLING |  |  |
|  | Demolishing brick work manually/ by mechanical means including stacking of serviceable material and disposal of unserviceable material within 50 metres lead as per direction of Engineer-in-charge. In cement mortar | 4 | M.CUBE |
|  | "Excavate foundation in all kinds of soil including soft rock except hard rock requiring blasting part return and fill in foundation and basement in layers not exceeding 15cm in depth including watering , breaking clods and consolidating, disposing and spreading surplus earth within the site upto 1.5m depth.(No extra shall be payable for working space) Tenderer to ensure proper dewatering and shoring methods to ensure all time dry condition. Any damages of sides of foundation or collapsing of soil from sides due to excessive water seepage or surface runoff or any other cause have to be replaced with quarry dust or Malaimann at no extra cost." | 10 | M.CUBE |
|  | Dismantling doors, windows and clerestory windows (steel or wood) shutter including chowkhats, architrave, holdfasts etc. complete and stacking within 50 metres lead : Of area beyond 3 sq. metres | 1 | each |
|  | Dismantling tile work in floors and roofs laid in cement mortar including stacking material within 50 metres lead. For thickness of tiles above 25 mm and up to 40 mm | 78 | M.SQ |
|  | Dismantling aluminium/ Gypsum partitions, doors, windows, fixed glazing and false ceiling including disposal of unserviceable material and stacking of serviceable material with in 50 meters lead as directed by Engineer-in-charge. | 60 | M.SQ |
|  | MODULAR CEILING PANELS - 50MM thick & Aluminium Ceiling Grid |  |  |
|  | Supply installation testing & commissioning (SITC) of the Modular False Ceiling System shall comprise of ceiling panels of double skin type 50mm thick, sandwiched with selfextinguishing quality PUF of density not less than 38kg / cum the inner and outer skin shall be of powder coated steel sheet, the thickness of inner skin and outer skin of the panel shall not less than 0.8mm the powder coated steel sheet shall be hot dip galvanized with zinc coating not less than 180gm/sqm both side of the GI sheet used. The false ceiling system shall also include ceiling grid system consisting of T-grid of extruded aluminum anodized surface, wall angle, suspender rod, turnbuckle, nuts, bolts, anchor fasteners etc., as required for fixing the panels, & HEPA filter frame, ceiling grid configuration length and width shall be suitable for fixing ceiling blanks, HEPA module, light fixtures etc. All the suspended material shall be galvanized finish. The ceiling grid system shall be suitable for walkable type false ceiling.the vertical panels to be from false ceiling till the true concrete ceiling. Fire rating - class O and colored approved by HLL. IMPORTANT NOTE: The contractor to take utmost care for tranport,shifting and erecting complete with all necessary hardware, accessories, fittings and fixtures, as per description in individual items, technical specifications, drawings and instructions of the Engineer-Incharge.Note:Necessary cutopens to be made for service outlets and finished after passing the service outlets. | 128 | M.SQ |
|  | DOORS |  |  |
|  | Providing laminated veneer lumber conforming to IS:14616 and TAD -15: 2001 (Part B) in factory made frames of doors, windows, clerestory windows and other frames, wrought framed and fixing in position with hold fast lugs or with dash fasteners of required dia & length (hold fast lugs or dash fastener shall be paid for separately). | QRO | M.CUBE |
|  | Providing and fixing 35 mm thick factory made laminated veneer lumber door shutter conforming to IS : 14616 and TADS 15:2001 (Part B), fixing with butt hinges of required size with necessary screws, all complete as per directions of Engineer- in-charge and panelling with panels of : (Note:- Butt hinges and necessary screws shall be paid separately) 12 mm thick one side Pre-laminated particle board (decorative lamination on one side and other sides balancing lamination) grade -1, medium density flat pressed, three layer particle board FPT - I or graded wood particle board FPT-1 conforming to IS : 3087 bonded with BWP type synthetic resin adhesive as per IS : 848 and pre-laminated conforming to IS : 12823, Grade -1, Type II marked : | QRO | M.SQ |
|  | STEELWORK |  |  |
|  | "Providing and fixing in position Guardian general purpose door manufactured by shakti met-Door limited or equivalent confirming to IS 3614(PART-2)1992 BS 4769(part-20&22) and ISO 834 or latest revisions with door frames of 1.20mm thick galvanised steel sheet press formed to rebated profile of size 100x57mm with chamfered bottom frame filled with concret and fixed in position with metallic expansion shield with counter sunk screws and door leaves made up of 0.80mm thick galvanised steel sheet press formed to provide a 46 mm thick fully flush,double skin door shell with lock seam joints at style edges withno visible joints ,screws,welding on the face of the shutter.Internal reinforcements provided at top,bottom and style edges for fire rating ,double shutter doors to be provided with a meeting styles.The rate is inclusive of necessary welding with columns. Door to be provided with rectangular/circular view panel as indicated in architectural drawings all as per manufacturer's specifications including providing necessary hardware like stainless steel imported steel butt hinges,handle ,mortice lock,door closure,handle from Dorma or equivalent make,handles all as per european standards, door stopper, painting with Duco spray paint standard ,painting finish of approved colour and shade over approved primer all complete.(Measurement of door for payment will be from the outer to outer frame in elevation only).Rate shall include cost of grouting with concrete in the frame and also welded in steel column/ frame. The rate should also includes cost of Silicon sealent at the junction of Wall/frame to be done on both sides." | 3.00 | M.SQ |
|  | Extra for providing and fixing auto closing device at bottom in either single leaf or double leaf shutter of Shaktimet door or equivalent including necessary hardwares etc., complete. Each shutter shall have one camlock assembly all as per manufacturer's specification and as directed.Camlock assembly components shall have colour similar to door colour. | 1 | EACH |
|  | Supplying fabricating and fixing in position insert plates of required size and shape in RCC members at all heights including necessary straightening ,cutting, holing, welding, lugs of required size and painting with a shop coat of anti corrosive primer paint all complete as directed. | 0.1 | M.T |
|  | """Supplying and fixing rolling shutters of 'Gandhi' or equivalent make made of 18 gauge, 80mm wide rolled MS lathe interlocked throughout their entire length and jointed together by end locks and mounted on specially designed pipe shaft with superior flet spring (enclosed on suitable boxes) complete with brackets, side guides and arrangements for inside and outside locking with mechanical devices for operation.The rate should include cost of providing necessary holes in the Masonry/RCC with pipe sleeve of suitable dia to facilate both side operation. Rolling Shutter to be designed to withstand a wind speed of 180KM/hr. (Only structural opening will be measured for the purpose of payment.)""" | 18 | M.SQ |
|  | Supply and Laying in roofing bare GALVALUME profiled sheet made out of 0.47mm TCT 550 Mpa cold rolled sheet with hot dip metallic coating of aluminium zinc alloy 150gms/sqm having 975/1000mm cover width with 26mm high crest at 195/200mm c/c in lengths as approved by Architect also to suit to site requirements and fixed with necessary suitable imported galvanized carbon steel self drilling & self tapping screws of BUILDEX or corroshield confirming to ASTM 3566-1990 (class- III) of required diameter and length as approved all complete with EPDM sealing washers, foam fillers wherever necessary, butyl sealent at all laps, silicon sealent wherever necessary. Rate should include cost of plastic caps of approved colour of U.V resistance and button bolts all as per manufacturer's specification all complete as approved by the Architect.(Only laid area without overlapping wil be measured for payment.). | 14 | M.SQ |
|  | Providing and fixing M.S. Tubular frames for doors, windows, ventilators and cupboard with rectangular/ L-Type sections, made of 1.60 mm thick M.S. Sheet, joints mitred, welded and grinded finish, with profiles of required size, including fixing of necessary butt hinges and screws and applying a priming coat of approved steel primer. Fixing with 15x3 mm lugs 10 cm long embedded in cement concrete block 15x10x10 cm of C.C. 1:3:6 (1 Cement : 3 coarse sand : 6 graded stone aggregate 20 mm nominal size) | 220 | KG |
|  | FLOORING & COVING |  |  |
|  | Providing and fixing 18 mm thick gang saw cut, mirror polished, premoulded and prepolished, machine cut for kitchen platforms, vanity counters, window sills, facias and similar locations of required size, approved shade, colour and texture laid over 20 mm thick base cement mortar 1:4 (1 cement : 4 coarse sand), joints treated with white cement, mixed with matching pigment, epoxy touch ups, including rubbing, curing, moulding and polishing to edges to give high gloss finish etc. complete at all levels.Granite of any colour and shade Area of slab over 0.50 sqm | 5 | M.SQ |
|  | Providing and laying vitrified floor tiles in different sizes (thickness to be specified by the manufacturer) with water absorption less than 0.08% and conforming to IS: 15622, of approved make, in all colours and shades, laid on 20mm thick cement mortar 1:4 (1 cement : 4 coarse sand), jointing with grey cement slurry @ 3.3 kg/ sqm including grouting the joints with white cement and matching pigments etc., complete. Size of Tile 600x600 mm | 42 | M.SQ |
|  | Providing and laying Vitrified tiles in different sizes (thickness to be specified by the manufacturer), with water absorption less than 0.08% and conforming to IS: 15622, of approved brand & manufacturer, in all colours and shade, in skirting, riser of steps, laid with cement based high polymer modified quick set tile adhesive (water based) conforming to IS: 15477, in average 6 mm thickness, including grouting of joints (Payment for grouting of joints to be made separately). | 68 | M.SQ |
|  | Grouting the joints of flooring tiles having joints of 3 mm width, using epoxy grout mix of 0.70 kg of organic coated filler of desired shade (0.10 kg of hardener and 0.20 kg of resin per kg), including filling / grouting and finishing complete as per direction of Engineer-in-charge. Size of Tile 600x600 mm | QRO | M.SQ |
|  | Supply installation testing & commissioning (SITC) of PVC coving D 50 with required aluminium / PVC back supports with application of clear sealant to the lips of the coving, complete with all necessary hardware, accessories, fittings and fixtures, as per description in individual items, technical specifications, drawings ans instructions of the Engineer-Incharge | 12 | M.SQ |
|  | Providing and laying 13 mm thick antistatic/conductive vinyl flooring of make Marvel Vinyl/ Electra plus or equivqlent of approved quality and colour conforming to the IS: 3462/ 1986 on the finished concrete or terrazzo tiles flooring or any hard surface with necessary approved adhesive for laying antistatic vinyl sheet/roll including cleaning the floor before and after laying of Viinyl flooring and providing thin copper strip of size about 20.00 x 0.20 mm along the periphery of room and below each joint of two sheets wherever directed with copper screws/ rivets @600 mm c/c .Two ends of copper strips (terminals) to be fixed to wall about 300 mm height for earthing of antistatic flooring all complete as per the direction of Engineer-in-charge.( NOTE any deduction beyond 0.5 Sqm will be deducted from the measurement) | 96 | M.SQ |
|  | MASONRY WORK |  |  |
|  | Brick work with common burnt clay F.P.S. (non modular) bricks of class designation 7.5 in superstructure above plinth level up to floor V level in all shapes and sizes in : Cement mortar 1:4 (1 cement : 4 coarse sand) | 5 | M.CUBE |
|  | Half brick masonry with common burnt clay F.P.S. (non modular) bricks of class designation 7.5 in superstructure above plinth level up to floor V level. Cement mortar 1:3 (1 cement :3 coarse sand) | QRO | M.SQ |
|  | Extra for providing and placing in position 2 Nos 6mm dia. M.S. bars at every third course of half brick masonry. | QRO | M.SQ |
|  | Cement concrete 1:4:8 (1 cement ,4 sand , 8 hard blue granite stone metal 40mm and down gauge in layers not exceeding 15 cms in depth well consolidated in foundation and levelling course under floors including necessary shuttering, in steps etc all complete as directed. | 4 | M.CUBE |
|  | "RCC M30 (design mix) using hard blue granite metal 20mm and down gauge including centering, shuttering, curing etc,complete but excluding cost of reinforcement. Note: Polythene sheet shall be provided under the slabs." In Columns in Foundation, Plinth Beam and Basement | 3 | M.CUBE |
|  | "High Yield Strength Deformed bars in all RCC items & floorings General Note: Sizes of RCC elements shown in the drawing are only indicative and subject to change .No extra shall be payable on this account. Any change in size of concrete elements resulting in change of area of shuttering due to change in quantity of concrete shall also not be payable"" | 300 | KG |
|  | 12 mm cement plaster finished with a floating coat of neat cement of mix : 1:3 (1 cement: 3 fine sand) | 86 | M.SQ |
|  | Surface Preparation Cleaning of surface free from dust. Dirt, moisture, oil & all foreign material from the floor. For floor cleaning all mechanized equipments are to be used like scrubber, sanders, hot air blowers, vacuum cleaners, moisture test machine etc. Once cleaned, the surface is roughened to prepare the surface for application of epoxy primer. | 66 | M.SQ |
|  | Levelling with epoxy screed Application of epoxy primer over prepared surface and then filling of cracks, crumbled joints small size damages in the epoxy floor with epoxy mortar having high compressive strength. The epoxy mortar should be so applied so as to acheve level of – 1 mm from adjoining top level of epoxy floor surface once the epoxy mortaris applied , the area should be left to dry for minimum 24 hours before final coat of EPU paint is applied | 0.5 | M.CUBE |
|  | EPU Coating – Providing and application of concrete primer on top of epoxy screed. This is followed by self Levelling PU top coating of specified shade and colour of 1 mm thick with self leveling epoxy material by means of special spike rollers and other tools so as to remove all trapped air bubbles and achieve smooth, glossy and leveled surface. | 66 | M.SQ |
| PAINTING |  |  |  |
|  | Prepare surfaces and paint two coats of plastic emulsion paint of approved colour over two coats of Gypsum punning of 6 mm minimum thick and primer all complete as per manufacturers specifications to all ceilings in all floors. | QRO | M.SQ |
|  | Finishing with Deluxe Multi surface paint system for interiors and exteriors using Primer as per manufacturers specifications : Two or more coats applied on walls @ 1.25 ltr/10 sqm over and including one coat of Special primer applied @ 0.75 ltr /10 sqm | 488.00 | M.SQ |
|  | Painting wood work with Deluxe Multi Surface Paint of required shade. Two or more coat applied @ 0.90 ltr/10 sqm over an under coat of primer applied @0.75 ltr/10 sqm of approved brand and manufacture | QRO | M.SQ |
|  | Painting Steel work with Deluxe Multi Surface Paint to give an even shade. Two or more coat applied @ 0.90 ltr/ 10 sqm over an under coat of primer applied @ 0.80 ltr/ 10 sqm of approved brand and manufacture | QRO | M.SQ |
|  | Providing and applying plaster of paris putty of 2 mm thickness over plastered surface to prepare the surface even and smooth complete. | 488.00 | M.SQ |
|  | Removing white or colour wash by scrapping and sand papering and preparing the surface smooth including necessary repairs to scratches etc. complete | 488.00 | M.SQ |
|  | 1.Two range of curtain cloth – Type 1 - II. |  |  |
|  | 2. Both for doors, windows and partions. |  |  |
|  | 3. Assorted colours plain & prints |  |  |
|  | 4.Stitched as per site/ size requirement |  |  |
|  | 5.Minor modifications in stitching as per site/ size requirement |  |  |
|  | 6.Preferred colour (Brown, Cream, Beige) |  |  |
|  | 7.Curtain price (including cloth & curtain rod) to be quoted in Squre metre |  |  |
|  | 8.Stitching for each curtain to be absorb in the price of the cloth generally 2 meters. |  |  |
|  | 9.Fast colours, water wash |  |  |
|  | 10.Fabric density Non Transparent Composition |  |  |
|  | 11.Fabric of curtain shall be soft texture, great shading, drape effect and nice dobby. |  |  |
|  | 12.The bracket shall be side mounted and front mounted as required |  |  |
|  | 13.First rod mount is 3.5 inches from the wall. |  |  |
|  | 14.Where curtain used as a partition/ long length, sufficient support from ceiling should be provided to avoid sagging. |  |  |
|  | 15.Rate quoted for fully stitched Curtain with top nepha with punch holes after lining bucramreinforced with SS / PVC islets and lower nepha 200mm double turned. |  |  |
|  | Type -I (Surgical Curtain) |  |  |
|  | Curtain fabric weight shall be up to 275-300 (±5%) gsm Composition of fabric shall be polyester yarn I.E. 300 denier fast colour sink age shall be (±5%) Weft and warp shall be polyester / cotton polyester mix | 0 | M.SQ |
|  | VERTICAL BLINDS |  |  |
|  | Providing and fixing of vertical/roller blinds width 100 mm regular quantity made from 100 % polyster/Fibre yarn coated with fade resistant polymer to ensure complete opaquety with zero light transmittance, having suitable coating on exterior, including all fixtures and fittings complete in all respect as per design and instruction of Engineer in-charge and as per site conditions. The colour will be approved by the competent authority Make-MAC/VISTA/NOVA or any equivalent make. | 6 | M.SQ |
| ALUMINUM WORK |  |  |  |
|  | For shutters of doors, windows & ventilators including providing and fixing hinges/ pivots and making provision for fixing of fittings wherever required including the cost of EPDM rubber / neoprene gasket required (Fittings shall be paid for separately) Powder coated aluminium (minimum thickness of powder coating 50 micron) |  | M.T |
|  | Providing and fixing aluminium work for doors, windows, ventilators and partitions with extruded built up standard tubular sections/appropriate Z sections and other sections of approved make conforming to IS: 733 and IS: 1285, fixing with dash fasteners of required dia and size, including necessary filling up the gaps atjunctions, i.e. at top, bottom and sides with required EPDM rubber/ neoprene gasket etc. Aluminium sections shall be smooth, rust free, straight, mitred and jointed mechanically wherever required including cleat angle, Aluminium snap beading for glazing / paneling, C.P. brass / stainless steel screws, all complete as per architectural drawings and the directions of Engineer-in-charge : Powder coated aluminium (minimum thickness of powder coating 50 micron) | 0.25 | MT |
|  | Providing and fixing bright finished brass 100 mm mortice latch and lock with 6 levers and a pair of lever handles of approved quality with necessary screws etc. complete. | 4.00 | EACH |
|  | Providing and fixing bright finished brass tower bolts (barrel type) with necessary screws etc. complete : 250x10 mm | 4.00 | EACH |
|  | Providing and fixing aluminium die cast body tubular type universal hydraulic door closer (having brand logo with ISI, IS : 3564, embossed on the body, door weight upto 35 kg and door width upto 700 mm), with necessary accessories and screws etc. complete. | 4.00 | EACH |
|  | Providing and fixing glazing in aluminum door, window, ventilator shutters and partitions etc. with EPDM rubber / neoprene gasket etc. complete as per the architectural drawings and the directions of Engineer-in- Charge. (Cost of aluminum snap beading shall be paid in basic item): With float glass panes of 8 mm thickness. | 18.00 | SQM |
|  | Providing and fixing anodised aluminium grill (anodised transparent or dyed to required shade according to IS: 1868 with minimum anodic coating of grade AC 15) of approved design/pattern, with approved standard section and fixed to the existing window frame with C.P. brass/ stainless steel screws @ 200 mm centre to centre, including cutting the grill to proper opening size for fixing and operation of handles and fixing approved anodised aluminium standard section around the opening, all complete as per requirement and direction of Engineer-in-charge. (Only weight of grill to be measured for payment). | QRO | M.T |
|  | Providing and fixing aluminum tubular handle bar 32 mm outer dia, 3.0 mm thick & 2100 mm long with SS screws etc. complete as per direction of Engineerin-Charge. Anodized (AC 15 ) aluminum tubular handle bar | 8.00 | EACH |
|  | Providing and fixing bright finished brass hanging type floor door stopper with necessary screws, etc. complete. | 4.00 | EACH |
|  | Providing and fixing double action hydraulic floor spring of approved brand and manufacture IS : 6315 marked, for doors including cost of cutting floors as required, embedding in floors and cover plates with brass pivot and single piece M.S. sheet outer box with slide plate etc. complete as per the direction of Engineer-in-charge. With Brass cover Plate | QRO | EACH |
|  | Filling the gap in between aluminium/ stone/ wood frame and adjacent RCC/ Brick/ Stone/ wood/ Ceramic/ Gypsum work by providing weather/structural non sag elastomeric PU sealant over backer rod of approved quality as per architectural drawings and direction of Engineer-in-charge complete, complying to ASTM C920, DIN 18540-F & ISO 11600. Upto 20 mm depth and 20 mm width | 24.00 | RM |
|  | Providing and fixing 12 mm thick pre laminated particle board flat pressed three layer or graded wood particle board conforming to IS: 12823 Grade I Type II, in paneling fixed in aluminum doors, windows shutters and partition frames with C.P. brass / stainless steel screws etc. complete as per architectural drawings and directions of Engineer-inCharge. Pre-laminated particle board with decorative lamination on both sides | 12.00 | SQM |
| PLUMBING |  |  |  |
|  | Providing and fixing wash basin with C.I. brackets, 15 mm dia CP Brass single hole basin mixer of approved quality and make, including painting of fittings and brackets, cutting and making good the walls wherever required:- (a) White Vitreous China Wash basin size 550x400 mm with a 15 mm CP Brass single hole basin mixer | 1 | EACH |
|  | Providing and fixing Stainless Steel A ISI 304 (18/8) kitchen sink as per IS:13983 with C.I. brackets and stainless steel plug 40 mm, including painting of fittings and brackets, cutting and making good the walls wherever required : Kitchen sink with drain board 510x1040 mm bowl depth 250 mm | 2 | EACH |
|  | Providing and fixing PVC Waste pipe for sink or wash basin including PVC waste fittings complete.:Semi rigid pipe 32mm dia | 2 | EACH |
|  | Providing and fixing CP Waste coupling for wash basin and sink of approved quality, colour and make. Waste coupling 31mm of 79mm length and 62mm breadth, weighing not less than 45 gms. Make:Hindware,Jaquar etc.., | 2 | EACH |
|  | Providing and fixing trap of self cleansing design with screwed down or hinged grating with or without vent arm complete, including cost of cutting and making good the walls and floors : 100 mm inlet and 100 mm outlet Sand cast iron S&S as per IS: 3989 | QRO | EACH |
|  | Extra for providing opening of required size & shape for wash basin/ kitchen sink in kitchen platform, vanity counter and similar location in marble/ Granite/stone work, including necessary holes for pillar taps etc. including moulding, rubbing and polishing of cut edges etc. complete. | 2 | EACH |
|  | Providing and fixing 600x450 mm beveled edge mirror of superior glass (of approved quality) complete with 6 mm thick hard board ground fixed to wooden cleats with C.P. brass screws and washers complete. | QRO | EACH |
|  | Providing and fixing PTMT towel rail complete with brackets fixed to wooden cleats with CP brass screws with concealed fitting arrangement of approved quality and colour. 600mm long towel rail with total length of 645mm, width 78mm and effective height of 88mm, weighing not less than 190gms. | QRO | EACH |
|  | Providing and fixing PTMT liquid soap container 109mm wide, 125mm high and 112mm distance from wall of standard shape with bracket of the same materials with snap fittings of approved quality and colour. weighing not less than 105 gms. | QRO | EACH |
|  | Providing and fixing floor mounted, white vitreous china single piece, double traps syphonic water closet of approved brand/make, shape, size and pattern including integrated white vitreous china cistern of capacity 10 litres with dual flushing system, including all fittings and fixtures with seat cover, cistern fittings, nuts, bolts and gasket etc including making connection with the existing P/S trap, complete in all respect as per directions of Engineer-in-Charge. | QRO | EACH |
|  | Providing & fixing white vitreous china water less urinal of size 600 x 330 x 315 mm having antibacterial /germs free ceramic surface, fixed with cartridge having debris catcher and hygiene seal. | QRO | EACH |
|  | Providing and fixing 3 layer PP-R (Poly propylene Random copolymer) pipes confirming to IS 15801, UV stabilized & anti - microbial fusion welded, having thermal stability for hot & cold water supply, including all PP - R plain & brass threaded polypropylene random fittings, i/c fixing the pipe with clamps at 1.00 m spacing. This includes testing of joints complete as per direction of Engineer-in-Charge. |  |  |
|  | PN - 16 Pipe, 25mm OD (SDR-7.4) | 20 | RM |
|  | PN - 10 Pipe, 50mm OD (SDR-11) | 20 | RM |
|  | Providing and fixing PPR Unions |  |  |
|  | PPR Union 25mm | 6 | EACH |
|  | PPR Union 50mm | 6 | EACH |
|  | Providing and fixing ball valve (brass) of approved quality, High or low pressure, with plastic floats complete : 25 mm nominal bore | 1 | EACH |
|  | Providing and fixing of Single Tap Scrub Station as per the specification as follows 1.Deep sloping basin 2. One High mount spout/taps with elbow operated flow regulator (length of flow regulator approximately 18-24cms) 3. Tap should have provision for manual aeration of water 4. 304 stainless steel polished and ground 5. Wall Mounted Design with provision for back panel wall bracket fixation 6. Provision for Scrub Solution Stand with long handle for pumping 7. Sizes a. Width – Approximately 90-100 cms b. Height of Back Panel of sink – 90 -95 cms c. Height of Sink from Ground – 100-110 cms d. Depth of Sink from Top of the Sink – 40-45 cms e. Height of Tap from the top of the sink – 25-28 cms f. Protrusion of Sink from Wall – 45-50 cms | 0 | EACH |
|  | Providing and fixing Elbow Tap of approved quality conforming to IS standards and weighing not less than 810 gms. 15 mm nominal bore | QRO | EACH |
|  | FURNITURES |  |  |
|  | Providing & Fixing mid back chair .The seat & back are made up of 12mm thick hotpressed plywood, upholstered with fabric & moulded with Polyurethane foam seat and back covers. The back foam should designed with contouredlumbar support for extra comfort .The seat has extra thick foam on front edge to give extra comfort.Mid back size 475 mm Width x 580 mm Height mm. And seat size 470 mm Width x 480 mm Thepolyurethane foam should bemoulded with density = 45 kg/m3 and hardness = 16 +/2 kg on hampdem machine at 25% compression .Armrest are injection moulded from black copolymer polypropylene. The mechanism should befixed 360 degree revolving type,Upright position locking, tilt tension adjustment, seat back tilting ratio 1:3. The pneumatic height adjustment has an adjustment stroke of 12 cm. The pedestal should be injection moulded in black 33 % glass filled in nylon 66 with 5 nos twin wheel castors. The pedestal should be 66.3 cm pitch-center diameter. | 0 | EACH |
|  | Supplying and fixing of Three Seater Waiting Area Visitor Chairs · Non-Magnetic SS perforated sheet seat and back shell made of high quality steel makes it strong and durable. · 1608mm (L)X 695 (H) X 555 (D). Legs and Armrests are made of cold rolled steel and are chrome plated for long life. Armrest and legs are anti rust and anti scratch. | 0 | EACH |
|  | Providing and Supplying table of size 2400W x 750 D x 750H, whose work surface is made of 18mm thick. Pre Laminated Particle Board (PLB) All work surface edges are duly sealed with 2mm thick PVC Edge banding. The modesty panel is made of 18mm thick. Pre Laminated Particle Board (PLB) All work surface edges are duly sealed with 2mm thick PVC Edge banding. The rectangular frame is fabricated component in 1.2mm thick CRCA (IS:513), Finish: Powder coated (Epoxy polyester). The leg is Fabricated component in 38mm x25mm x1.2mm thick MS ERW Tube (IS:7138), Finish: Powder coat (Epoxy polyester).The leveler glide for leg is Nylon 6 & MS Bolt. The shell and drawer tray are made of 0.6mm thick CRCA (IS:513), Finish: Powder coat (Epoxy polyester).The drawer front is 0.8mm thick CRCA (IS:513), Finish: Powder coat (Epoxy polyester) and frame assembly is 1.2mm thick CRCA (IS:513), Finish: Powder coat (Epoxy polyester). | QRO | EACH |
|  | Providing and Supplying shoe rack of size 1175W x 1175 H x 375 D, whose work surface is made of 25mm thick. Pre Laminated Particle Board (PLB) OR Engineered Wood All work surface edges are duly sealed with 2mm thick PVC Edge banding. The modesty panel is made of 18mm thick. Pre Laminated Particle Board (PLB) All work surface edges are duly sealed with 2mm thick PVC Edge banding. Powder coat (Epoxy polyester). Brand: Evok OR Equivelent Assembly: Carpenter Assembly Dimensions: 1175W x 1175 H x 375 D Primary Material: Engineered Wood Room Type: Living Room | 1 | EACH |
|  | Providing & Fixing Sofa-2 seater. Size: 1320 mm (W) x 835 mm (D) x 690 mm (H) Seat Height from gound 415 mm. Frame is made up of Tropical Solid Wood with syntheatic leather Upholstery,Seat Foam : Polyurethane Foam with recron layer.Foam Density : 18 kg/ cubic cm. Back Foam : Polyurethane Foam with silicon fill at arms. Foam Density : 22 kg/cubic cm S-Spring webbing.The Sofa is based on wooden legs | QRO | EACH |
|  | Providing & Fixing HIGH back executive chair .The seat & back are made up of 12mm thick hotpressed plywood, upholstered with fabric & moulded with Polyurethane foam seat and back covers. The back foam should designed with contouredlumbar support for extra comfort .The seat has extra thick foam on front edge to give extra comfort.Mid back size 475 mm Width x 580 mm Height mm. And seat size 470 mm Width x 480 mm Thepolyurethane foam should bemoulded with density = 45 kg/m3 and hardness = 16 +/2 kg on hampdem machine at 25% compression .Armrest are injection moulded from black copolymer polypropylene. The mechanism should befixed 360 degree revolving type,Upright position locking, tilt tension adjustment, seat back tilting ratio 1:3. The pneumatic height adjustment has an adjustment stroke of 12 cm. The pedestal should be injection moulded in black 33 % glass filled in nylon 66 with 5 nos twin wheel castors. The pedestal should be 66.3 cm pitch-center diameter. | QRO | EACH |
|  | Supply of the FEATHERLITE Versaline Executive Table as per the Featherlite specifications . 1.V-MT 1890, W 1800 x D 900 x H 750 2. V-ST 1245 (LH & RH), W 1200 x D 450 x H 750 and 3. V-CZ, W 1200 x D 450 x H 750 | QRO | EACH |
|  | Conference Room table: Fabricated at site Conference System U shaped, open ended 20 seating in size of 1350x600x740 for two seats and 675x600x740 for one seat capacity with Table Top made of 25mm thick MDF with waterfall trop edge finished in pvc memberane using unique wrap around for better aesthetics & long lasting use. Understructure - Legs are made of 25mm thick standard black matt colour post formed panels connected with 18mm thick std black prelaminated modesty panels. Wire management - conecaled wire management to take care both horizontally & vertically. For open configurations, wire channel is mounted on the inner side of modesty panel for horizontal wire carrying. Vertical wire management is through flexible wire manager. Std table top height is 740mm, pvc membrane tops areseamless surface, post formed legs for safe use. | QRO | EACH |
|  | ELECTRICAL |  |  |
|  | Point wiring in PVC conduit with modular type switch. |  |  |
|  | Wiring for light point/fan point/exhaust fan point /call bell point with 1.5 sq mm FR PVC insulated copper conductor single core cable in surface/recessed medium class PVC conduit, with modular switch, modular plate, suitable PVC box and earthing the point with 1.5 sq.mm FR PVC insulated copper conductor single core cable etc as required. |  |  |
|  | Group C | 42 | Point |
|  | Wiring through PVC conduit |  |  |
|  | Wiring for circuit/submain wiring along with earth wire with the following sizes of FR PVC insulated copper conductor, single core cable in surface/recessed medium class PVC conduit as required. |  |  |
|  | 2X1.5 sq.mm + 1X1.5 sq.mm earth wire | 215 | Metre |
|  | 2X2.5 sq.mm + 1 X2.5 sq mm earth wire | 160 | Metre |
|  | 2X4 sq.mm + 1X4 sq.mm earth wire | 160 | Metre |
|  | A/C point |  |  |
|  | S/F 20A SPN MCB industrial socket outlet: Supplying and fixing 20 amps, 240 volts, SPN industrial type, socket outlet, with 2 pole and earth, metal enclosed plug top alongwith 20 amps "C" curve, SP, MCB, in sheet steel enclosure, on surface or in recess, with chained metal cover for the socket out let and complete with connections, testing and commissioning etc. as required. | 2 | Point |
|  | Modular type blanking plate: |  |  |
|  | Supplying and fixing modular blanking plate on the existing modular plate & switch box excluding modular plate as required. | 2 | Each |
|  | Modular boxes, bases & cover plate: |  |  |
|  | Supplying and fixing following size/modules, PVC box along with modular base & cover plate for modular switches in recess etc as required. |  |  |
|  | 2 module | 2 | Each |
|  | 3 module | 12 | Each |
|  | 6 module | 6 | Each |
|  | 8 module | 42 | Each |
|  | Modular type switch/socket: |  |  |
|  | Supplying and fixing following modular switch/socket on the existing modular plate & switch box including connections but excluding modular plate etc as required. |  |  |
|  | 5/6 Amps switch | 12 | Each |
|  | 3 pin 5/6 Amp socket outlet | 12 | Each |
|  | 15/16 Amps switch | 42 | Each |
|  | 6pin 15/16 Amps (universal) socket outlet | 42 | Each |
|  | Telephone socket outlet | 4 | Each |
|  | Telephone wiring in existing conduit: |  |  |
|  | Supplying and drawing following pair 0.5 sq mm FR PVC insulated annealed copper conductor, unarmored telephone cable in the existing surface/ recessed steel/ PVC conduit as required. |  |  |
|  | 2 Pair | 100 | Metre |
|  | S/F PVC conduit: |  |  |
|  | Supplying and fixing of following sizes of medium class PVC conduit along with accessories in surface/recess including cutting the wall and making good the same in case of recessed conduit as required. |  |  |
|  | 20 mm | 100 | Metre |
|  | 32 mm | 100 | Metre |
|  | ITC Fluorescent fitting directly on surface: Installation, testing and commissioning of pre-wired, fluorescent fitting / compact fluorescent fitting of all types, complete with all accessories and tube etc. directly on ceiling/ wall, including connection with 1.5 sq. mm FR PVC insulated, copper conductor, single core cable and earthing etc. as required. | 12 | Each |
|  | ITC Ceiling fan: Installation, testing and commissioning of ceiling fan, including wiring the down rods of standard length (upto 30 cm) with 1.5 sq. mm FR PVC insulated, copper conductor, single core cable etc. as required. | 2 | Each |
|  | CABLE LAYING |  |  |
|  | Laying and fixing of one number PVC insulated and PVC sheathed / XLPE power cable of 1.1kV grade of following size clamped on wall surface with 1mm thick saddles as required complete. |  |  |
|  | From VDB to LDB |  |  |
|  | 4X16 sq.mm | 100 | Metre |
|  | Supplying and drawing of UTP 4 pair CAT 6 LAN Cable in the existing surface/ recessed Steel/ PVC conduit as required. 2 run of cable | 0 | RM |
|  | Supply and installation of 20W recess mounted, decorative Soft lite luminaire, powder coated CRCA housing complete with all electrical accessories like electronic ballast, lamp holder, 20W LED etc complete with all mounting accessories like chain, anchor fastner etc as required. (Make Crompton,Havells) | 12 | Each |
|  | Supplying and fixing PVC angle batten/batten holder with 12W LED, including connection as required. | 4 | Each |
|  | Supply of 1x28W flourescent tube fitting sleek with decorative end cap, powder coated CRCA steel channel complete with all electrical accessories like electronic ballast, lamp holder, 28W T5 lamp etc complete as required. (Philips TMS 540 HF / Wipro WRF 80128 SG / GE DMS 28 N N) | 0 | Each |
|  | Supply & installation of 1x9 watts recess mounted down light fixture, with aluminium reflector, CRCA sheet housing, complete with all accessories like ballast, 9W LED lamp etc complete as required. | 0 | Each |
|  | Supply and installation of Ceiling fan with its all accessories(Make Havells/bajaj/usha) | 2 | Each |
|  | Miscellaneous Items |  |  |
|  | Air Conditioner Supply, installation & Commissioning of 2 ton capacity split type air conditioner (3 star rating ) - Voltas / LG / Carrier / Blue star with suitable stabiliser ( V-Guard / Everest / Voltas / Microteck), Out Door unit stand , 4mtr draining pipe.necessery copper piping .1 Year warranty for AC & Stabilizer 5 Years warranty for compressor , Bidder should mention the make and model of the Air Conditioner and stabilizer | 2 | EACH |
|  | Supply and pointing exhaust fan shutter type white colour-9"/10"(Make Havells/Bajaj/Almonard) | 0 | Each |
|  | Supply & Fixing of calling bell Intercom system with all accessories. [ Bazer ] | 0 | Each |

Section – VII

Quality Control Requirements

(Proforma for equipment and quality control employed by the manufacturer(s)

Tender Reference No.

Date of opening

Time

Name and address of the Tenderer:

Note:  All the following details shall relate to the manufacturer(s) for the goods quoted for.

01	Name of the manufacturer

a.	Full postal address with e mail address.

b.	Telephone number

c.	Fax number

d.        E-mail

02	 Quality control arrangement details

a.	For final product evaluation

03	Test certificate held

a	. Type test

b	. BIS/ISO certification

c	. Any other

Signature and seal of the Tenderer

Section – VIII

Qualification Criteria

| Sl. No. | Minimum Eligibility Criteria | Narrations |
| --- | --- | --- |
| 1 | Qualification of Tenderer | The Tenderer must be a Manufacturer or Authorized Agent/dealer/Distributor or Consortium partner. |
| 2 | Entirety | The tenderer shall quote for all items in the list of requirements, if any of the item not quoted then their bid will be treated as non responsive. |
| 3 | Experience | The tenderer shall have successfully established atleast One Microbiology Section on turnkey basis during last three years prior the date of Tender opening with similar equipment performing similar functions and meeting major specification parameters of the quoted item (s), which is functioning satisfactorily in India. |
| 4 | Financial capability | The annual sales turnover for the bidder should be minimum of Rs. 1 Crore in any of the last 3 financial years: During the year (2016-2017 to 2018-2019), and it should be certified by the chartered accountants. by Chartered Accountant shall be enclosed– Proforma ‘A’- I |
| 5 | Responsiveness to List of Equipments | The items quoted shall be responsive to the requirements in the List of Equipments in the Tender. |

Note

- In support of 3, the Tenderer shall furnish Performance statement in the enclosed Proforma ‘A’  II (SECTION- IX) duly signed by the Purchaser

- The Tenderer shall furnish a brief write-up, packed with adequate data explaining and establishing his available capacity/capability (both technical and financial) to perform the Contract (if awarded) within the stipulated time period, after meeting all its current/present commitments. The Tenderer shall also furnish details of Equipment and Quality Control in the enclosed Section VII.

- Not withstanding anything stated above, the Purchaser reserves the right to assess the Tenderer’s capability and capacity to perform the contract  satisfactorily before deciding on award of Contract, should circumstances warrant such an assessment in the overall interest of the Purchaser.

- The Purchaser reserves the right to ask for a free demonstration of the quoted equipment at a pre determined place acceptable to the purchaser for technical acceptability as per the tender specifications, before the opening of the Price Tender.

###### Section – IX A

TENDER FORM

Date__________

To

Deputy General Manager (Materials),

HLL Lifecare Limited,

Akkulam Factory, Sreekariam PO,

Thiruvananthapuram – 17.

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: hcdcmo@lifecarehll.com

TENDER No. _________dated ___________

We, the undersigned have examined the above mentioned TE document, including amendment/corrigendum No. __________, dated ________ (if any), the receipt of which is hereby confirmed. We now offer to supply and deliver___________ (Description of goods and services) in conformity with your above referred document for the sum of _____________ (total tender amount in figures and words), as shown in the price schedule(s), attached herewith and made part of this tender.

If our tender is accepted, we undertake to supply the goods and perform the services as mentioned above, in accordance with the delivery schedule specified in the List of Requirements.

We further confirm that, if our tender is accepted, we shall provide you with a performance security of required amount in an acceptable form in terms of GCC clause 5, read with modification, if any, in Section - IV – “Special Conditions of Contract”, for due performance of the contract.

We agree to keep our tender valid for acceptance as required in the GIT clause 20, or for subsequently extended period, if any, agreed to by us. We also accordingly confirm to abide by this tender up to the aforesaid period and this tender may be accepted any time before the expiry of the aforesaid period. We further confirm that, until a formal contract is executed, this tender read with your written acceptance thereof within the aforesaid period shall constitute a binding contract between us.

We further understand that you are not bound to accept the lowest or any tender you may receive against your above-referred tender enquiry.

We confirm that we do not stand deregistered/banned/blacklisted by any Govt. Authorities.

We confirm that we fully agree to the terms and conditions specified in above mentioned TE document, including amendment/ corrigendum if any

(Signature with date)

(Name and designation) Duly authorised to sign tender for and on behalf of

###### PROFORMA ‘A’ I

Name & Address of

Tenderer		:

Whether Manufacturer

or agent		:

------------------------------------------------------------------------------------------------------------------------------

Financial Data	16-17               17-18            18-1       Average annual turnover for last 3 years

------------------------------------------------------------------------------------------------------------------------------

Annual Turn over

Gross Profit/Loss

Whether attached the relevant

Pages of Balance sheet and

Profit and Loss A/c duly certified

By Chartered Accountant

Signature and stamp of Chartered Accountant		 Signature of Tenderer with stamp

###### PROFORMA ‘A’ II  (Documentary proof for Proforma ‘A’)

# PROFORMA FOR PERFORMANCE STATEMENT

(For the period of last three years)

Tender Reference No.	: _________________________________

Date of opening	: _________________________________

Order cross  reference No.	: _________________________________

# Name and address of Purchaser	: _________________________________

Country of origin,Name and address

of the manufacturer/Tenderer	: _________________________________

| Order placed by (full address of Purchaser | Order number and date | Description and quantity of ordered goods and services | Value of order (Rs.) | Date of completion of Contract |  | Remarks indicating reasons for delay if any | Have the goods been functioning Satisfactorily |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | As per contract | Actual |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |

Date:

Signature and seal of the Purchaser

NB: Satisfactory performance certificate from clients to be enclosed

### Section – IXB

TENDER FORM (for price bid)

Date__________

To

Deputy General Manager (Materials),

HLL Lifecare Limited,

Akkulam Factory, Sreekariam PO,

Thiruvananthapuram – 17.

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: hcdcmo@lifecarehll.com

Tender No. _________dated ___________

We, the undersigned have examined the above mentioned TE document, including amendment/corrigendum No. __________, dated ________ (if any), the receipt of which is hereby confirmed. We now offer to supply and deliver___________ (Description of goods and services) in conformity with your above referred document for the sum of _____________ (total tender amount in figures and words), as shown in the price schedule(s), attached herewith and made part of this tender.

If our tender is accepted, we undertake to supply the goods and perform the services as mentioned above, in accordance with the delivery schedule specified in the List of Requirements.

We further confirm that, if our tender is accepted, we shall provide you with a performance security of required amount in an acceptable form in terms of GCC clause 5, read with modification, if any, in Section - IV – “Special Conditions of Contract”, for due performance of the contract.

We agree to keep our tender valid for acceptance as required in the GIT clause 20, read with modification, if any in – “Special Instructions to Tenderers” or for subsequently extended period, if any, agreed to by us. We also accordingly confirm to abide by this tender up to the aforesaid period and this tender may be accepted any time before the expiry of the aforesaid period. We further confirm that, until a formal contract is executed, this tender read with your written acceptance thereof within the aforesaid period shall constitute a binding contract between us.

We further understand that you are not bound to accept the lowest or any tender you may receive against your above-referred tender enquiry.

We confirm that we do not stand deregistered/banned/blacklisted by any Govt. Authorities.

We confirm that we fully agree to the terms and conditions specified in above mentioned TE document, including amendment/ corrigendum if any

(Signature with date)

(Name and designation) Duly authorised to sign tender for and on behalf of

SECTION – X     PRICE SCHEDULE (A)

- A)	SUPPLY OF EQUIPMENTS TO MICROBIOLOGY SECTION TO THE STATE DRUG TESTING LABORATORY, TRIPURA

| 1 | 2 | 3 | 4 |  | 5 |  |  |  |  |  |  |  | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SI N o | Brief Description of Goods | Brand / Model | Make / Country of Origin | Qty | Price per unit (Rs.) |  |  |  |  |  |  |  | Total Price (at Consignee Site) basis (Rs.) 4 x 5(f) |
|  |  |  |  |  | Basic Rate Per Unit | GST |  | Freight& Insurance & Delivery at Destination |  | Service charges, if any, specify | Other charges, if any, specify | Unit Price (at Consignee Site) bas |  |
|  |  |  |  |  |  |  |  |  |  |  |  | (f) =a+b+c+d+e |  |
|  |  |  |  |  | (a) | (b) |  | € |  | (d) | € |  |  |
|  |  |  |  |  |  | % | Amt | % | Amt |  |  |  |  |
| 1 | AHU |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 2 | Anaerobic Gas .Jar |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 3 | Antibiotic zone scale |  |  | 3 |  |  |  |  |  |  |  |  |  |
| 4 | Autoclave Horizontal (79 L) |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 5 | Autoclave Vertical |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 6 | Automated microbial detection system |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 7 | Beaker 10 ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 8 | Beaker 50 ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 9 | Beaker 250 ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 10 | Beaker 500 ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 11 | Beaker 1000 ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 12 | Bio safety cabinet Class III-CE certified |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 13 | BOD Incubator 500 |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 14 | BOD Incubator 600 |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 15 | BOD Incubator 900 |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 16 | Centrifuge Refrigerated |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 17 | Electronic Balance (Micro/Semi Micro) |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 18 | Computer Unit |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 19 | Conical Flask 250ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 20 | Conical Flask 500ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 21 | Conical Flask 1000ml |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 22 | Cross Over Bench (Stainless Steel) |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 23 | Digital Colony Counter |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 24 | Dropping Bottle tbr Stains (125 ml) |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 25 | Electric Loop sterilizer |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 26 | Flask flat Bottom (1000 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 27 | Fumi gation apparatus |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 28 | Garment Shelf (Stainless Steel) |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 29 | Heating Block |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 30 | Hot Air Oven (Digitally Control) |  |  | 3 |  |  |  |  |  |  |  |  |  |
| 31 | Hygrometer |  |  | 4 |  |  |  |  |  |  |  |  |  |
| 32 | Incubators |  |  | 4 |  |  |  |  |  |  |  |  |  |
| 33 | Induction Cook Top |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 34 | Lanrinar Air flow cabinet |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 35 | Measuring Cylinder (10 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 36 | Measuring Cylinder (25 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 37 | Measuring Cylinder (50 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 38 | Measuring Cylinder (100 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 39 | Measuring Cylinder (250 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 40 | Measuring Cylinder (500 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 41 | Measuring Cylinder (1000 ml) |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 42 | Membrane Filtration Unit With all Accessories |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 43 | Micrometer Stage |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 44 | Microscope |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 45 | Micrometer Eye Piece |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 46 | Microwave Oven |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 47 | Multichannel Pipette-variable volume 5- 50µl, 20-200µl |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 48 | Magnetic Stirrer with Bead |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 49 | Petridish with lid 50x17mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 50 | Petridish with lid 80x17mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 51 | Petridish with lid 100x17 mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 52 | Petridish with lid 50x12 mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 53 | Petridish with lid 80x 15mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 54 | Petridish with lid 90x 15mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 55 | Petridish with lid 100x 15mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 56 | Petridish with lid 150x 25mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 57 | Petridish with lid 200x 30mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 58 | pH Detern'rination apparatus |  |  | 3 |  |  |  |  |  |  |  |  |  |
| 59 | Volumetric Pipette 1 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 60 | Volumetric Pipette 2 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 61 | Volumetric Pipette 5 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 62 | Volumetric Pipette 10 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 63 | Volumetric Pipette 20 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 64 | Volumetric Pipette 25 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 65 | Pipette Graduated 1 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 66 | Pipette Graduated 2 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 67 | Pipette Graduated 5 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 68 | Pipette Graduated 10 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 69 | Pipette Graduated 25 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 70 | Pipette Sterilization Can |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 71 | Reagent Bottles/ Stopper - 100 cc |  |  | 60 |  |  |  |  |  |  |  |  |  |
| 72 | Reagent Bottles/ Stopper - 1000 cc |  |  | 36 |  |  |  |  |  |  |  |  |  |
| 73 | Reagent Bottles/ Stopper - 250 cc |  |  | 24 |  |  |  |  |  |  |  |  |  |
| 74 | Reagent Bottles/ Stopper - 50 cc |  |  | 60 |  |  |  |  |  |  |  |  |  |
| 75 | Refrigerator (300 ltr or More) |  |  | 3 |  |  |  |  |  |  |  |  |  |
| 76 | Single channel Pipette-variable volume 0.5-l0µl, 5-50µl, 50-200µl,200-1000 µl |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 77 | Slides |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 78 | Staining (troug) rod (10 slides) |  |  | 60 |  |  |  |  |  |  |  |  |  |
| 79 | Stainless Steel Bead (1000pcs per box) |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 80 | Table (Stainless Steel) |  |  | 3 |  |  |  |  |  |  |  |  |  |
| 81 | Test tube l00 x l2 mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 82 | Test tube l50 x l8 mnr |  |  | 10 |  |  |  |  |  |  |  |  |  |
| 83 | Test tube 75 x 12 mm |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 84 | Thermometer |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 85 | Stool (Stainless Steel) |  |  | 5 |  |  |  |  |  |  |  |  |  |
| 86 | Trinocular Fluorescent microscope with digital colour CD Camera LED |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 87 | Ultrapure Water Purifier |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 88 | UPS (3 kV) |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 89 | UV Larnp |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 90 | Voluntetric Flask 50 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 91 | Voluntetric Flask 100 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 92 | Voluntetric Flask 250 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 93 | Voluntetric Flask 500 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 94 | Voluntetric Flask 1000 ml |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 95 | Vortex Mixer |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 96 | Water bath |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 97 | Water bath serological 37 deg |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 98 | Water bath serological 56 deg |  |  | 1 |  |  |  |  |  |  |  |  |  |
| 99 | Micropipette Tips (variable Volume) 0-200 µl |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 100 | Micropipette Tips (variable Volume) 200-1000 µl |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 101 | Micropipette Tips (variable Volume) 0-10 µl |  |  | 20 |  |  |  |  |  |  |  |  |  |
| 102 | Ultra Sonic bath |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 103 | Vial Decapping machine |  |  | 2 |  |  |  |  |  |  |  |  |  |
| 104 | Civil & Electrical works |  |  | 1 |  |  |  |  |  |  |  |  |  |
|  | Total |  |  |  |  |  |  |  |  |  |  |  |  |

Total tender price in rupees________________________________________________________________

In words: _____________________________________________________________________________________________________

- Note:

- If there is a discrepancy between the unit price and total price THE UNIT PRICE shall prevail.

- Successful  tenderer shall execute the contract with HLL as per Section XIV.

- The charges for CAMC after warranty shall be quoted separately as per Section – X – Price Schedule (B)

Name________________________

Business Address________________________

Place: __________________________                                                                                Signature of Tenderer________________________

Date: _________________________

SECTION – X     PRICE SCHEDULE (A)

- A-1)	PRICE SCHEDULE FOR TURNKEY WORKS

| S.No | DESCRIPTION | QTY | UNIT | Unit Rate in Rs. | Total Rate in Rs. |
| --- | --- | --- | --- | --- | --- |
|  | DISMANTLING |  |  |  |  |
|  | Demolishing brick work manually/ by mechanical means including stacking of serviceable material and disposal of unserviceable material within 50 metres lead as per direction of Engineer-in-charge. In cement mortar | 4 | M.CUBE |  |  |
|  | "Excavate foundation in all kinds of soil including soft rock except hard rock requiring blasting part return and fill in foundation and basement in layers not exceeding 15cm in depth including watering , breaking clods and consolidating, disposing and spreading surplus earth within the site upto 1.5m depth.(No extra shall be payable for working space) Tenderer to ensure proper dewatering and shoring methods to ensure all time dry condition. Any damages of sides of foundation or collapsing of soil from sides due to excessive water seepage or surface runoff or any other cause have to be replaced with quarry dust or Malaimann at no extra cost." | 10 | M.CUBE |  |  |
|  | Dismantling doors, windows and clerestory windows (steel or wood) shutter including chowkhats, architrave, holdfasts etc. complete and stacking within 50 metres lead : Of area beyond 3 sq. metres | 1 | each |  |  |
|  | Dismantling tile work in floors and roofs laid in cement mortar including stacking material within 50 metres lead. For thickness of tiles above 25 mm and up to 40 mm | 78 | M.SQ |  |  |
|  | Dismantling aluminium/ Gypsum partitions, doors, windows, fixed glazing and false ceiling including disposal of unserviceable material and stacking of serviceable material with in 50 meters lead as directed by Engineer-in-charge. | 60 | M.SQ |  |  |
|  | MODULAR CEILING PANELS - 50MM thick & Aluminium Ceiling Grid |  |  |  |  |
|  | Supply installation testing & commissioning (SITC) of the Modular False Ceiling System shall comprise of ceiling panels of double skin type 50mm thick, sandwiched with selfextinguishing quality PUF of density not less than 38kg / cum the inner and outer skin shall be of powder coated steel sheet, the thickness of inner skin and outer skin of the panel shall not less than 0.8mm the powder coated steel sheet shall be hot dip galvanized with zinc coating not less than 180gm/sqm both side of the GI sheet used. The false ceiling system shall also include ceiling grid system consisting of T-grid of extruded aluminum anodized surface, wall angle, suspender rod, turnbuckle, nuts, bolts, anchor fasteners etc., as required for fixing the panels, & HEPA filter frame, ceiling grid configuration length and width shall be suitable for fixing ceiling blanks, HEPA module, light fixtures etc. All the suspended material shall be galvanized finish. The ceiling grid system shall be suitable for walkable type false ceiling.the vertical panels to be from false ceiling till the true concrete ceiling. Fire rating - class O and colored approved by HLL. IMPORTANT NOTE: The contractor to take utmost care for tranport,shifting and erecting complete with all necessary hardware, accessories, fittings and fixtures, as per description in individual items, technical specifications, drawings and instructions of the Engineer-Incharge.Note:Necessary cutopens to be made for service outlets and finished after passing the service outlets. | 128 | M.SQ |  |  |
|  | DOORS |  |  |  |  |
|  | Providing laminated veneer lumber conforming to IS:14616 and TAD -15: 2001 (Part B) in factory made frames of doors, windows, clerestory windows and other frames, wrought framed and fixing in position with hold fast lugs or with dash fasteners of required dia & length (hold fast lugs or dash fastener shall be paid for separately). | QRO | M.CUBE |  |  |
|  | Providing and fixing 35 mm thick factory made laminated veneer lumber door shutter conforming to IS : 14616 and TADS 15:2001 (Part B), fixing with butt hinges of required size with necessary screws, all complete as per directions of Engineer- in-charge and panelling with panels of : (Note:- Butt hinges and necessary screws shall be paid separately) 12 mm thick one side Pre-laminated particle board (decorative lamination on one side and other sides balancing lamination) grade -1, medium density flat pressed, three layer particle board FPT - I or graded wood particle board FPT-1 conforming to IS : 3087 bonded with BWP type synthetic resin adhesive as per IS : 848 and pre-laminated conforming to IS : 12823, Grade -1, Type II marked : | QRO | M.SQ |  |  |
|  | STEELWORK |  |  |  |  |
|  | "Providing and fixing in position Guardian general purpose door manufactured by shakti met-Door limited or equivalent confirming to IS 3614(PART-2)1992 BS 4769(part-20&22) and ISO 834 or latest revisions with door frames of 1.20mm thick galvanised steel sheet press formed to rebated profile of size 100x57mm with chamfered bottom frame filled with concret and fixed in position with metallic expansion shield with counter sunk screws and door leaves made up of 0.80mm thick galvanised steel sheet press formed to provide a 46 mm thick fully flush,double skin door shell with lock seam joints at style edges withno visible joints ,screws,welding on the face of the shutter.Internal reinforcements provided at top,bottom and style edges for fire rating ,double shutter doors to be provided with a meeting styles.The rate is inclusive of necessary welding with columns. Door to be provided with rectangular/circular view panel as indicated in architectural drawings all as per manufacturer's specifications including providing necessary hardware like stainless steel imported steel butt hinges,handle ,mortice lock,door closure,handle from Dorma or equivalent make,handles all as per european standards, door stopper, painting with Duco spray paint standard ,painting finish of approved colour and shade over approved primer all complete.(Measurement of door for payment will be from the outer to outer frame in elevation only).Rate shall include cost of grouting with concrete in the frame and also welded in steel column/ frame. The rate should also includes cost of Silicon sealent at the junction of Wall/frame to be done on both sides." | 3.00 | M.SQ |  |  |
|  | Extra for providing and fixing auto closing device at bottom in either single leaf or double leaf shutter of Shaktimet door or equivalent including necessary hardwares etc., complete. Each shutter shall have one camlock assembly all as per manufacturer's specification and as directed.Camlock assembly components shall have colour similar to door colour. | 1 | EACH |  |  |
|  | Supplying fabricating and fixing in position insert plates of required size and shape in RCC members at all heights including necessary straightening ,cutting, holing, welding, lugs of required size and painting with a shop coat of anti corrosive primer paint all complete as directed. | 0.1 | M.T |  |  |
|  | """Supplying and fixing rolling shutters of 'Gandhi' or equivalent make made of 18 gauge, 80mm wide rolled MS lathe interlocked throughout their entire length and jointed together by end locks and mounted on specially designed pipe shaft with superior flet spring (enclosed on suitable boxes) complete with brackets, side guides and arrangements for inside and outside locking with mechanical devices for operation.The rate should include cost of providing necessary holes in the Masonry/RCC with pipe sleeve of suitable dia to facilate both side operation. Rolling Shutter to be designed to withstand a wind speed of 180KM/hr. (Only structural opening will be measured for the purpose of payment.)""" | 18 | M.SQ |  |  |
|  | Supply and Laying in roofing bare GALVALUME profiled sheet made out of 0.47mm TCT 550 Mpa cold rolled sheet with hot dip metallic coating of aluminium zinc alloy 150gms/sqm having 975/1000mm cover width with 26mm high crest at 195/200mm c/c in lengths as approved by Architect also to suit to site requirements and fixed with necessary suitable imported galvanized carbon steel self drilling & self tapping screws of BUILDEX or corroshield confirming to ASTM 3566-1990 (class- III) of required diameter and length as approved all complete with EPDM sealing washers, foam fillers wherever necessary, butyl sealent at all laps, silicon sealent wherever necessary. Rate should include cost of plastic caps of approved colour of U.V resistance and button bolts all as per manufacturer's specification all complete as approved by the Architect.(Only laid area without overlapping wil be measured for payment.). | 14 | M.SQ |  |  |
|  | Providing and fixing M.S. Tubular frames for doors, windows, ventilators and cupboard with rectangular/ L-Type sections, made of 1.60 mm thick M.S. Sheet, joints mitred, welded and grinded finish, with profiles of required size, including fixing of necessary butt hinges and screws and applying a priming coat of approved steel primer. Fixing with 15x3 mm lugs 10 cm long embedded in cement concrete block 15x10x10 cm of C.C. 1:3:6 (1 Cement : 3 coarse sand : 6 graded stone aggregate 20 mm nominal size) | 220 | KG |  |  |
|  | FLOORING & COVING |  |  |  |  |
|  | Providing and fixing 18 mm thick gang saw cut, mirror polished, premoulded and prepolished, machine cut for kitchen platforms, vanity counters, window sills, facias and similar locations of required size, approved shade, colour and texture laid over 20 mm thick base cement mortar 1:4 (1 cement : 4 coarse sand), joints treated with white cement, mixed with matching pigment, epoxy touch ups, including rubbing, curing, moulding and polishing to edges to give high gloss finish etc. complete at all levels.Granite of any colour and shade Area of slab over 0.50 sqm | 5 | M.SQ |  |  |
|  | Providing and laying vitrified floor tiles in different sizes (thickness to be specified by the manufacturer) with water absorption less than 0.08% and conforming to IS: 15622, of approved make, in all colours and shades, laid on 20mm thick cement mortar 1:4 (1 cement : 4 coarse sand), jointing with grey cement slurry @ 3.3 kg/ sqm including grouting the joints with white cement and matching pigments etc., complete. Size of Tile 600x600 mm | 42 | M.SQ |  |  |
|  | Providing and laying Vitrified tiles in different sizes (thickness to be specified by the manufacturer), with water absorption less than 0.08% and conforming to IS: 15622, of approved brand & manufacturer, in all colours and shade, in skirting, riser of steps, laid with cement based high polymer modified quick set tile adhesive (water based) conforming to IS: 15477, in average 6 mm thickness, including grouting of joints (Payment for grouting of joints to be made separately). | 68 | M.SQ |  |  |
|  | Grouting the joints of flooring tiles having joints of 3 mm width, using epoxy grout mix of 0.70 kg of organic coated filler of desired shade (0.10 kg of hardener and 0.20 kg of resin per kg), including filling / grouting and finishing complete as per direction of Engineer-in-charge. Size of Tile 600x600 mm | QRO | M.SQ |  |  |
|  | Supply installation testing & commissioning (SITC) of PVC coving D 50 with required aluminium / PVC back supports with application of clear sealant to the lips of the coving, complete with all necessary hardware, accessories, fittings and fixtures, as per description in individual items, technical specifications, drawings ans instructions of the Engineer-Incharge | 12 | M.SQ |  |  |
|  | Providing and laying 13 mm thick antistatic/conductive vinyl flooring of make Marvel Vinyl/ Electra plus or equivqlent of approved quality and colour conforming to the IS: 3462/ 1986 on the finished concrete or terrazzo tiles flooring or any hard surface with necessary approved adhesive for laying antistatic vinyl sheet/roll including cleaning the floor before and after laying of Viinyl flooring and providing thin copper strip of size about 20.00 x 0.20 mm along the periphery of room and below each joint of two sheets wherever directed with copper screws/ rivets @600 mm c/c .Two ends of copper strips (terminals) to be fixed to wall about 300 mm height for earthing of antistatic flooring all complete as per the direction of Engineer-in-charge.( NOTE any deduction beyond 0.5 Sqm will be deducted from the measurement) | 96 | M.SQ |  |  |
|  | MASONRY WORK |  |  |  |  |
|  | Brick work with common burnt clay F.P.S. (non modular) bricks of class designation 7.5 in superstructure above plinth level up to floor V level in all shapes and sizes in : Cement mortar 1:4 (1 cement : 4 coarse sand) | 5 | M.CUBE |  |  |
|  | Half brick masonry with common burnt clay F.P.S. (non modular) bricks of class designation 7.5 in superstructure above plinth level up to floor V level. Cement mortar 1:3 (1 cement :3 coarse sand) | QRO | M.SQ |  |  |
|  | Extra for providing and placing in position 2 Nos 6mm dia. M.S. bars at every third course of half brick masonry. | QRO | M.SQ |  |  |
|  | Cement concrete 1:4:8 (1 cement ,4 sand , 8 hard blue granite stone metal 40mm and down gauge in layers not exceeding 15 cms in depth well consolidated in foundation and levelling course under floors including necessary shuttering, in steps etc all complete as directed. | 4 | M.CUBE |  |  |
|  | "RCC M30 (design mix) using hard blue granite metal 20mm and down gauge including centering, shuttering, curing etc,complete but excluding cost of reinforcement. Note: Polythene sheet shall be provided under the slabs." In Columns in Foundation, Plinth Beam and Basement | 3 | M.CUBE |  |  |
|  | "High Yield Strength Deformed bars in all RCC items & floorings General Note: Sizes of RCC elements shown in the drawing are only indicative and subject to change .No extra shall be payable on this account. Any change in size of concrete elements resulting in change of area of shuttering due to change in quantity of concrete shall also not be payable"" | 300 | KG |  |  |
|  | 12 mm cement plaster finished with a floating coat of neat cement of mix : 1:3 (1 cement: 3 fine sand) | 86 | M.SQ |  |  |
|  | Surface Preparation Cleaning of surface free from dust. Dirt, moisture, oil & all foreign material from the floor. For floor cleaning all mechanized equipments are to be used like scrubber, sanders, hot air blowers, vacuum cleaners, moisture test machine etc. Once cleaned, the surface is roughened to prepare the surface for application of epoxy primer. | 66 | M.SQ |  |  |
|  | Levelling with epoxy screed Application of epoxy primer over prepared surface and then filling of cracks, crumbled joints small size damages in the epoxy floor with epoxy mortar having high compressive strength. The epoxy mortar should be so applied so as to acheve level of – 1 mm from adjoining top level of epoxy floor surface once the epoxy mortaris applied , the area should be left to dry for minimum 24 hours before final coat of EPU paint is applied | 0.5 | M.CUBE |  |  |
|  | EPU Coating – Providing and application of concrete primer on top of epoxy screed. This is followed by self Levelling PU top coating of specified shade and colour of 1 mm thick with self leveling epoxy material by means of special spike rollers and other tools so as to remove all trapped air bubbles and achieve smooth, glossy and leveled surface. | 66 | M.SQ |  |  |
| PAINTING |  |  |  |  |  |
|  | Prepare surfaces and paint two coats of plastic emulsion paint of approved colour over two coats of Gypsum punning of 6 mm minimum thick and primer all complete as per manufacturers specifications to all ceilings in all floors. | QRO | M.SQ |  |  |
|  | Finishing with Deluxe Multi surface paint system for interiors and exteriors using Primer as per manufacturers specifications : Two or more coats applied on walls @ 1.25 ltr/10 sqm over and including one coat of Special primer applied @ 0.75 ltr /10 sqm | 488.00 | M.SQ |  |  |
|  | Painting wood work with Deluxe Multi Surface Paint of required shade. Two or more coat applied @ 0.90 ltr/10 sqm over an under coat of primer applied @0.75 ltr/10 sqm of approved brand and manufacture | QRO | M.SQ |  |  |
|  | Painting Steel work with Deluxe Multi Surface Paint to give an even shade. Two or more coat applied @ 0.90 ltr/ 10 sqm over an under coat of primer applied @ 0.80 ltr/ 10 sqm of approved brand and manufacture | QRO | M.SQ |  |  |
|  | Providing and applying plaster of paris putty of 2 mm thickness over plastered surface to prepare the surface even and smooth complete. | 488.00 | M.SQ |  |  |
|  | Removing white or colour wash by scrapping and sand papering and preparing the surface smooth including necessary repairs to scratches etc. complete | 488.00 | M.SQ |  |  |
|  | 1.Two range of curtain cloth – Type 1 - II. |  |  |  |  |
|  | 2. Both for doors, windows and partions. |  |  |  |  |
|  | 3. Assorted colours plain & prints |  |  |  |  |
|  | 4.Stitched as per site/ size requirement |  |  |  |  |
|  | 5.Minor modifications in stitching as per site/ size requirement |  |  |  |  |
|  | 6.Preferred colour (Brown, Cream, Beige) |  |  |  |  |
|  | 7.Curtain price (including cloth & curtain rod) to be quoted in Squre metre |  |  |  |  |
|  | 8.Stitching for each curtain to be absorb in the price of the cloth generally 2 meters. |  |  |  |  |
|  | 9.Fast colours, water wash |  |  |  |  |
|  | 10.Fabric density Non Transparent Composition |  |  |  |  |
|  | 11.Fabric of curtain shall be soft texture, great shading, drape effect and nice dobby. |  |  |  |  |
|  | 12.The bracket shall be side mounted and front mounted as required |  |  |  |  |
|  | 13.First rod mount is 3.5 inches from the wall. |  |  |  |  |
|  | 14.Where curtain used as a partition/ long length, sufficient support from ceiling should be provided to avoid sagging. |  |  |  |  |
|  | 15.Rate quoted for fully stitched Curtain with top nepha with punch holes after lining bucramreinforced with SS / PVC islets and lower nepha 200mm double turned. |  |  |  |  |
|  | Type -I (Surgical Curtain) |  |  |  |  |
|  | Curtain fabric weight shall be up to 275-300 (±5%) gsm Composition of fabric shall be polyester yarn I.E. 300 denier fast colour sink age shall be (±5%) Weft and warp shall be polyester / cotton polyester mix | 0 | M.SQ |  |  |
|  | VERTICAL BLINDS |  |  |  |  |
|  | Providing and fixing of vertical/roller blinds width 100 mm regular quantity made from 100 % polyster/Fibre yarn coated with fade resistant polymer to ensure complete opaquety with zero light transmittance, having suitable coating on exterior, including all fixtures and fittings complete in all respect as per design and instruction of Engineer in-charge and as per site conditions. The colour will be approved by the competent authority Make-MAC/VISTA/NOVA or any equivalent make. | 6 | M.SQ |  |  |
| ALUMINUM WORK |  |  |  |  |  |
|  | For shutters of doors, windows & ventilators including providing and fixing hinges/ pivots and making provision for fixing of fittings wherever required including the cost of EPDM rubber / neoprene gasket required (Fittings shall be paid for separately) Powder coated aluminium (minimum thickness of powder coating 50 micron) |  | M.T |  |  |
|  | Providing and fixing aluminium work for doors, windows, ventilators and partitions with extruded built up standard tubular sections/appropriate Z sections and other sections of approved make conforming to IS: 733 and IS: 1285, fixing with dash fasteners of required dia and size, including necessary filling up the gaps atjunctions, i.e. at top, bottom and sides with required EPDM rubber/ neoprene gasket etc. Aluminium sections shall be smooth, rust free, straight, mitred and jointed mechanically wherever required including cleat angle, Aluminium snap beading for glazing / paneling, C.P. brass / stainless steel screws, all complete as per architectural drawings and the directions of Engineer-in-charge : Powder coated aluminium (minimum thickness of powder coating 50 micron) | 0.25 | MT |  |  |
|  | Providing and fixing bright finished brass 100 mm mortice latch and lock with 6 levers and a pair of lever handles of approved quality with necessary screws etc. complete. | 4.00 | EACH |  |  |
|  | Providing and fixing bright finished brass tower bolts (barrel type) with necessary screws etc. complete : 250x10 mm | 4.00 | EACH |  |  |
|  | Providing and fixing aluminium die cast body tubular type universal hydraulic door closer (having brand logo with ISI, IS : 3564, embossed on the body, door weight upto 35 kg and door width upto 700 mm), with necessary accessories and screws etc. complete. | 4.00 | EACH |  |  |
|  | Providing and fixing glazing in aluminum door, window, ventilator shutters and partitions etc. with EPDM rubber / neoprene gasket etc. complete as per the architectural drawings and the directions of Engineer-in- Charge. (Cost of aluminum snap beading shall be paid in basic item): With float glass panes of 8 mm thickness. | 18.00 | SQM |  |  |
|  | Providing and fixing anodised aluminium grill (anodised transparent or dyed to required shade according to IS: 1868 with minimum anodic coating of grade AC 15) of approved design/pattern, with approved standard section and fixed to the existing window frame with C.P. brass/ stainless steel screws @ 200 mm centre to centre, including cutting the grill to proper opening size for fixing and operation of handles and fixing approved anodised aluminium standard section around the opening, all complete as per requirement and direction of Engineer-in-charge. (Only weight of grill to be measured for payment). | QRO | M.T |  |  |
|  | Providing and fixing aluminum tubular handle bar 32 mm outer dia, 3.0 mm thick & 2100 mm long with SS screws etc. complete as per direction of Engineerin-Charge. Anodized (AC 15 ) aluminum tubular handle bar | 8.00 | EACH |  |  |
|  | Providing and fixing bright finished brass hanging type floor door stopper with necessary screws, etc. complete. | 4.00 | EACH |  |  |
|  | Providing and fixing double action hydraulic floor spring of approved brand and manufacture IS : 6315 marked, for doors including cost of cutting floors as required, embedding in floors and cover plates with brass pivot and single piece M.S. sheet outer box with slide plate etc. complete as per the direction of Engineer-in-charge. With Brass cover Plate | QRO | EACH |  |  |
|  | Filling the gap in between aluminium/ stone/ wood frame and adjacent RCC/ Brick/ Stone/ wood/ Ceramic/ Gypsum work by providing weather/structural non sag elastomeric PU sealant over backer rod of approved quality as per architectural drawings and direction of Engineer-in-charge complete, complying to ASTM C920, DIN 18540-F & ISO 11600. Upto 20 mm depth and 20 mm width | 24.00 | RM |  |  |
|  | Providing and fixing 12 mm thick pre laminated particle board flat pressed three layer or graded wood particle board conforming to IS: 12823 Grade I Type II, in paneling fixed in aluminum doors, windows shutters and partition frames with C.P. brass / stainless steel screws etc. complete as per architectural drawings and directions of Engineer-inCharge. Pre-laminated particle board with decorative lamination on both sides | 12.00 | SQM |  |  |
| PLUMBING |  |  |  |  |  |
|  | Providing and fixing wash basin with C.I. brackets, 15 mm dia CP Brass single hole basin mixer of approved quality and make, including painting of fittings and brackets, cutting and making good the walls wherever required:- (a) White Vitreous China Wash basin size 550x400 mm with a 15 mm CP Brass single hole basin mixer | 1 | EACH |  |  |
|  | Providing and fixing Stainless Steel A ISI 304 (18/8) kitchen sink as per IS:13983 with C.I. brackets and stainless steel plug 40 mm, including painting of fittings and brackets, cutting and making good the walls wherever required : Kitchen sink with drain board 510x1040 mm bowl depth 250 mm | 2 | EACH |  |  |
|  | Providing and fixing PVC Waste pipe for sink or wash basin including PVC waste fittings complete.:Semi rigid pipe 32mm dia | 2 | EACH |  |  |
|  | Providing and fixing CP Waste coupling for wash basin and sink of approved quality, colour and make. Waste coupling 31mm of 79mm length and 62mm breadth, weighing not less than 45 gms. Make:Hindware,Jaquar etc.., | 2 | EACH |  |  |
|  | Providing and fixing trap of self cleansing design with screwed down or hinged grating with or without vent arm complete, including cost of cutting and making good the walls and floors : 100 mm inlet and 100 mm outlet Sand cast iron S&S as per IS: 3989 | QRO | EACH |  |  |
|  | Extra for providing opening of required size & shape for wash basin/ kitchen sink in kitchen platform, vanity counter and similar location in marble/ Granite/stone work, including necessary holes for pillar taps etc. including moulding, rubbing and polishing of cut edges etc. complete. | 2 | EACH |  |  |
|  | Providing and fixing 600x450 mm beveled edge mirror of superior glass (of approved quality) complete with 6 mm thick hard board ground fixed to wooden cleats with C.P. brass screws and washers complete. | QRO | EACH |  |  |
|  | Providing and fixing PTMT towel rail complete with brackets fixed to wooden cleats with CP brass screws with concealed fitting arrangement of approved quality and colour. 600mm long towel rail with total length of 645mm, width 78mm and effective height of 88mm, weighing not less than 190gms. | QRO | EACH |  |  |
|  | Providing and fixing PTMT liquid soap container 109mm wide, 125mm high and 112mm distance from wall of standard shape with bracket of the same materials with snap fittings of approved quality and colour. weighing not less than 105 gms. | QRO | EACH |  |  |
|  | Providing and fixing floor mounted, white vitreous china single piece, double traps syphonic water closet of approved brand/make, shape, size and pattern including integrated white vitreous china cistern of capacity 10 litres with dual flushing system, including all fittings and fixtures with seat cover, cistern fittings, nuts, bolts and gasket etc including making connection with the existing P/S trap, complete in all respect as per directions of Engineer-in-Charge. | QRO | EACH |  |  |
|  | Providing & fixing white vitreous china water less urinal of size 600 x 330 x 315 mm having antibacterial /germs free ceramic surface, fixed with cartridge having debris catcher and hygiene seal. | QRO | EACH |  |  |
|  | Providing and fixing 3 layer PP-R (Poly propylene Random copolymer) pipes confirming to IS 15801, UV stabilized & anti - microbial fusion welded, having thermal stability for hot & cold water supply, including all PP - R plain & brass threaded polypropylene random fittings, i/c fixing the pipe with clamps at 1.00 m spacing. This includes testing of joints complete as per direction of Engineer-in-Charge. |  |  |  |  |
|  | PN - 16 Pipe, 25mm OD (SDR-7.4) | 20 | RM |  |  |
|  | PN - 10 Pipe, 50mm OD (SDR-11) | 20 | RM |  |  |
|  | Providing and fixing PPR Unions |  |  |  |  |
|  | PPR Union 25mm | 6 | EACH |  |  |
|  | PPR Union 50mm | 6 | EACH |  |  |
|  | Providing and fixing ball valve (brass) of approved quality, High or low pressure, with plastic floats complete : 25 mm nominal bore | 1 | EACH |  |  |
|  | Providing and fixing of Single Tap Scrub Station as per the specification as follows 1.Deep sloping basin 2. One High mount spout/taps with elbow operated flow regulator (length of flow regulator approximately 18-24cms) 3. Tap should have provision for manual aeration of water 4. 304 stainless steel polished and ground 5. Wall Mounted Design with provision for back panel wall bracket fixation 6. Provision for Scrub Solution Stand with long handle for pumping 7. Sizes a. Width – Approximately 90-100 cms b. Height of Back Panel of sink – 90 -95 cms c. Height of Sink from Ground – 100-110 cms d. Depth of Sink from Top of the Sink – 40-45 cms e. Height of Tap from the top of the sink – 25-28 cms f. Protrusion of Sink from Wall – 45-50 cms | 0 | EACH |  |  |
|  | Providing and fixing Elbow Tap of approved quality conforming to IS standards and weighing not less than 810 gms. 15 mm nominal bore | QRO | EACH |  |  |
|  | FURNITURES |  |  |  |  |
|  | Providing & Fixing mid back chair .The seat & back are made up of 12mm thick hotpressed plywood, upholstered with fabric & moulded with Polyurethane foam seat and back covers. The back foam should designed with contouredlumbar support for extra comfort .The seat has extra thick foam on front edge to give extra comfort.Mid back size 475 mm Width x 580 mm Height mm. And seat size 470 mm Width x 480 mm Thepolyurethane foam should bemoulded with density = 45 kg/m3 and hardness = 16 +/2 kg on hampdem machine at 25% compression .Armrest are injection moulded from black copolymer polypropylene. The mechanism should befixed 360 degree revolving type,Upright position locking, tilt tension adjustment, seat back tilting ratio 1:3. The pneumatic height adjustment has an adjustment stroke of 12 cm. The pedestal should be injection moulded in black 33 % glass filled in nylon 66 with 5 nos twin wheel castors. The pedestal should be 66.3 cm pitch-center diameter. | 0 | EACH |  |  |
|  | Supplying and fixing of Three Seater Waiting Area Visitor Chairs · Non-Magnetic SS perforated sheet seat and back shell made of high quality steel makes it strong and durable. · 1608mm (L)X 695 (H) X 555 (D). Legs and Armrests are made of cold rolled steel and are chrome plated for long life. Armrest and legs are anti rust and anti scratch. | 0 | EACH |  |  |
|  | Providing and Supplying table of size 2400W x 750 D x 750H, whose work surface is made of 18mm thick. Pre Laminated Particle Board (PLB) All work surface edges are duly sealed with 2mm thick PVC Edge banding. The modesty panel is made of 18mm thick. Pre Laminated Particle Board (PLB) All work surface edges are duly sealed with 2mm thick PVC Edge banding. The rectangular frame is fabricated component in 1.2mm thick CRCA (IS:513), Finish: Powder coated (Epoxy polyester). The leg is Fabricated component in 38mm x25mm x1.2mm thick MS ERW Tube (IS:7138), Finish: Powder coat (Epoxy polyester).The leveler glide for leg is Nylon 6 & MS Bolt. The shell and drawer tray are made of 0.6mm thick CRCA (IS:513), Finish: Powder coat (Epoxy polyester).The drawer front is 0.8mm thick CRCA (IS:513), Finish: Powder coat (Epoxy polyester) and frame assembly is 1.2mm thick CRCA (IS:513), Finish: Powder coat (Epoxy polyester). | QRO | EACH |  |  |
|  | Providing and Supplying shoe rack of size 1175W x 1175 H x 375 D, whose work surface is made of 25mm thick. Pre Laminated Particle Board (PLB) OR Engineered Wood All work surface edges are duly sealed with 2mm thick PVC Edge banding. The modesty panel is made of 18mm thick. Pre Laminated Particle Board (PLB) All work surface edges are duly sealed with 2mm thick PVC Edge banding. Powder coat (Epoxy polyester). Brand: Evok OR Equivelent Assembly: Carpenter Assembly Dimensions: 1175W x 1175 H x 375 D Primary Material: Engineered Wood Room Type: Living Room | 1 | EACH |  |  |
|  | Providing & Fixing Sofa-2 seater. Size: 1320 mm (W) x 835 mm (D) x 690 mm (H) Seat Height from gound 415 mm. Frame is made up of Tropical Solid Wood with syntheatic leather Upholstery,Seat Foam : Polyurethane Foam with recron layer.Foam Density : 18 kg/ cubic cm. Back Foam : Polyurethane Foam with silicon fill at arms. Foam Density : 22 kg/cubic cm S-Spring webbing.The Sofa is based on wooden legs | QRO | EACH |  |  |
|  | Providing & Fixing HIGH back executive chair .The seat & back are made up of 12mm thick hotpressed plywood, upholstered with fabric & moulded with Polyurethane foam seat and back covers. The back foam should designed with contouredlumbar support for extra comfort .The seat has extra thick foam on front edge to give extra comfort.Mid back size 475 mm Width x 580 mm Height mm. And seat size 470 mm Width x 480 mm Thepolyurethane foam should bemoulded with density = 45 kg/m3 and hardness = 16 +/2 kg on hampdem machine at 25% compression .Armrest are injection moulded from black copolymer polypropylene. The mechanism should befixed 360 degree revolving type,Upright position locking, tilt tension adjustment, seat back tilting ratio 1:3. The pneumatic height adjustment has an adjustment stroke of 12 cm. The pedestal should be injection moulded in black 33 % glass filled in nylon 66 with 5 nos twin wheel castors. The pedestal should be 66.3 cm pitch-center diameter. | QRO | EACH |  |  |
|  | Supply of the FEATHERLITE Versaline Executive Table as per the Featherlite specifications . 1.V-MT 1890, W 1800 x D 900 x H 750 2. V-ST 1245 (LH & RH), W 1200 x D 450 x H 750 and 3. V-CZ, W 1200 x D 450 x H 750 | QRO | EACH |  |  |
|  | Conference Room table: Fabricated at site Conference System U shaped, open ended 20 seating in size of 1350x600x740 for two seats and 675x600x740 for one seat capacity with Table Top made of 25mm thick MDF with waterfall trop edge finished in pvc memberane using unique wrap around for better aesthetics & long lasting use. Understructure - Legs are made of 25mm thick standard black matt colour post formed panels connected with 18mm thick std black prelaminated modesty panels. Wire management - conecaled wire management to take care both horizontally & vertically. For open configurations, wire channel is mounted on the inner side of modesty panel for horizontal wire carrying. Vertical wire management is through flexible wire manager. Std table top height is 740mm, pvc membrane tops areseamless surface, post formed legs for safe use. | QRO | EACH |  |  |
|  | ELECTRICAL |  |  |  |  |
|  | Point wiring in PVC conduit with modular type switch. |  |  |  |  |
|  | Wiring for light point/fan point/exhaust fan point /call bell point with 1.5 sq mm FR PVC insulated copper conductor single core cable in surface/recessed medium class PVC conduit, with modular switch, modular plate, suitable PVC box and earthing the point with 1.5 sq.mm FR PVC insulated copper conductor single core cable etc as required. |  |  |  |  |
|  | Group C | 42 | Point |  |  |
|  | Wiring through PVC conduit |  |  |  |  |
|  | Wiring for circuit/submain wiring along with earth wire with the following sizes of FR PVC insulated copper conductor, single core cable in surface/recessed medium class PVC conduit as required. |  |  |  |  |
|  | 2X1.5 sq.mm + 1X1.5 sq.mm earth wire | 215 | Metre |  |  |
|  | 2X2.5 sq.mm + 1 X2.5 sq mm earth wire | 160 | Metre |  |  |
|  | 2X4 sq.mm + 1X4 sq.mm earth wire | 160 | Metre |  |  |
|  | A/C point |  |  |  |  |
|  | S/F 20A SPN MCB industrial socket outlet: Supplying and fixing 20 amps, 240 volts, SPN industrial type, socket outlet, with 2 pole and earth, metal enclosed plug top alongwith 20 amps "C" curve, SP, MCB, in sheet steel enclosure, on surface or in recess, with chained metal cover for the socket out let and complete with connections, testing and commissioning etc. as required. | 2 | Point |  |  |
|  | Modular type blanking plate: |  |  |  |  |
|  | Supplying and fixing modular blanking plate on the existing modular plate & switch box excluding modular plate as required. | 2 | Each |  |  |
|  | Modular boxes, bases & cover plate: |  |  |  |  |
|  | Supplying and fixing following size/modules, PVC box along with modular base & cover plate for modular switches in recess etc as required. |  |  |  |  |
|  | 2 module | 2 | Each |  |  |
|  | 3 module | 12 | Each |  |  |
|  | 6 module | 6 | Each |  |  |
|  | 8 module | 42 | Each |  |  |
|  | Modular type switch/socket: |  |  |  |  |
|  | Supplying and fixing following modular switch/socket on the existing modular plate & switch box including connections but excluding modular plate etc as required. |  |  |  |  |
|  | 5/6 Amps switch | 12 | Each |  |  |
|  | 3 pin 5/6 Amp socket outlet | 12 | Each |  |  |
|  | 15/16 Amps switch | 42 | Each |  |  |
|  | 6pin 15/16 Amps (universal) socket outlet | 42 | Each |  |  |
|  | Telephone socket outlet | 4 | Each |  |  |
|  | Telephone wiring in existing conduit: |  |  |  |  |
|  | Supplying and drawing following pair 0.5 sq mm FR PVC insulated annealed copper conductor, unarmored telephone cable in the existing surface/ recessed steel/ PVC conduit as required. |  |  |  |  |
|  | 2 Pair | 100 | Metre |  |  |
|  | S/F PVC conduit: |  |  |  |  |
|  | Supplying and fixing of following sizes of medium class PVC conduit along with accessories in surface/recess including cutting the wall and making good the same in case of recessed conduit as required. |  |  |  |  |
|  | 20 mm | 100 | Metre |  |  |
|  | 32 mm | 100 | Metre |  |  |
|  | ITC Fluorescent fitting directly on surface: Installation, testing and commissioning of pre-wired, fluorescent fitting / compact fluorescent fitting of all types, complete with all accessories and tube etc. directly on ceiling/ wall, including connection with 1.5 sq. mm FR PVC insulated, copper conductor, single core cable and earthing etc. as required. | 12 | Each |  |  |
|  | ITC Ceiling fan: Installation, testing and commissioning of ceiling fan, including wiring the down rods of standard length (upto 30 cm) with 1.5 sq. mm FR PVC insulated, copper conductor, single core cable etc. as required. | 2 | Each |  |  |
|  | CABLE LAYING |  |  |  |  |
|  | Laying and fixing of one number PVC insulated and PVC sheathed / XLPE power cable of 1.1kV grade of following size clamped on wall surface with 1mm thick saddles as required complete. |  |  |  |  |
|  | From VDB to LDB |  |  |  |  |
|  | 4X16 sq.mm | 100 | Metre |  |  |
|  | Supplying and drawing of UTP 4 pair CAT 6 LAN Cable in the existing surface/ recessed Steel/ PVC conduit as required. 2 run of cable | 0 | RM |  |  |
|  | Supply and installation of 20W recess mounted, decorative Soft lite luminaire, powder coated CRCA housing complete with all electrical accessories like electronic ballast, lamp holder, 20W LED etc complete with all mounting accessories like chain, anchor fastner etc as required. (Make Crompton,Havells) | 12 | Each |  |  |
|  | Supplying and fixing PVC angle batten/batten holder with 12W LED, including connection as required. | 4 | Each |  |  |
|  | Supply of 1x28W flourescent tube fitting sleek with decorative end cap, powder coated CRCA steel channel complete with all electrical accessories like electronic ballast, lamp holder, 28W T5 lamp etc complete as required. (Philips TMS 540 HF / Wipro WRF 80128 SG / GE DMS 28 N N) | 0 | Each |  |  |
|  | Supply & installation of 1x9 watts recess mounted down light fixture, with aluminium reflector, CRCA sheet housing, complete with all accessories like ballast, 9W LED lamp etc complete as required. | 0 | Each |  |  |
|  | Supply and installation of Ceiling fan with its all accessories(Make Havells/bajaj/usha) | 2 | Each |  |  |
|  | Miscellaneous Items |  |  |  |  |
|  | Air Conditioner Supply, installation & Commissioning of 2 ton capacity split type air conditioner (3 star rating ) - Voltas / LG / Carrier / Blue star with suitable stabiliser ( V-Guard / Everest / Voltas / Microteck), Out Door unit stand , 4mtr draining pipe.necessery copper piping .1 Year warranty for AC & Stabilizer 5 Years warranty for compressor , Bidder should mention the make and model of the Air Conditioner and stabilizer | 2 | EACH |  |  |
|  | Supply and pointing exhaust fan shutter type white colour-9"/10"(Make Havells/Bajaj/Almonard) | 0 | Each |  |  |
|  | Supply & Fixing of calling bell Intercom system with all accessories. [ Bazer ] | 0 | Each |  |  |
|  | Sub Total |  |  |  |  |
|  | GST @ 12% |  |  |  |  |
|  | Grand Total |  |  |  |  |

CIVIL DRAWING- EXISTING SITE

CIVIL DRAWING- PROPOSED SITE

SECTION – X PRICE SCHEDULE (B)

- PRICE SCHEDULE FOR COMPREHENSIVE ANNUAL MAINTENANCE CONTRACT AFTER WARRANTY PERIOD

| 1 | 2 | 3 | 4 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| SI No. | Brief Description of Goods | QTY | Comprehensive Maintenance Contract price per unit per year |  |  |  |  |
|  |  |  | 1st year | 2nd year | 3rd year | 4th year | 5th year |

* After completion of 5 years Warranty period for respective equipment as mentioned in the warranty clause

NOTE:-

1. In case of discrepancy between unit price and total prices, THE UNIT PRICE shall prevail.

2. The cost of CAMC which includes preventive maintenance including testing & calibration as per technical/ service /operational manual, labour and spares, after satisfactory completion of Warranty period may be quoted for next 5 years on yearly basis for complete equipment and Turnkey.

3. The cost of CAMC may be quoted along with taxes applicable on the date of Tender Opening. The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such taxes and no claim for the same will be entertained later.

4. Successful tenderer shall enter into an conrtact with THE STATE DRUG TESTING LABORATORY, TRIPURA, after expiration of warranty period with approval from HLL as per Section XV.

5. The payment of CAMC will be made as per clause GCC clause 21.1.

6. The uptime warranty will be 95 % on 24 (hrs) X 7 (days) X 365 (days) basis or as stated in Technical Specification of the TE document.

7.  All software updates should be provided free of cost during CAMC period.

8. The stipulations in Technical Specification will supersede above provisions.

9. The supplier shall keep sufficient stock of spares required during Comprehensive Annual Maintenance Contract period. In case the spares are required to be   imported, it would be the responsibility of the supplier to import and get them custom cleared and pay all necessary duties.

Name________________________

Business Address________________________

Place: ___________________________                             Signature of Tenderer________________________

Date: _________________________                                  Seal of the Tenderer________________________

SECTION – X PRICE SCHEDULE (C)

- PRICE SCHEDULE FOR CONSUMABLES NOT COVERED UNDER WARRANTY & CAMC

| 1 | 2 | 3 | 4 |  | 5 |  |  |  |  |  |  |  | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SI No | Brief Description of Goods | Brand / Model | Make / Country of Origin | Qty | Price per unit (Rs.) |  |  |  |  |  |  |  | Total Price (at Consignee Site) basis (Rs.) 4 x 5(f) |
|  |  |  |  |  | Basic Rate Per Unit | GST |  | Freight& Insurance & Delivery at Destination |  | Service charges, if any, specify | Other charges, if any, specify | Unit Price (at Consignee Site) bas |  |
|  |  |  |  |  |  |  |  |  |  |  |  | (f) =a+b+c+d+e |  |
|  |  |  |  |  | (a) | (b) |  | € |  | (d) | € |  |  |
|  |  |  |  |  |  | % | Amt | % | Amt |  |  |  |  |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |  |  |  |

*

Name________________________

Business Address________________________

Place: ___________________________                             Signature of Tenderer________________________

Date: _________________________                                  Seal of the Tenderer________________________

SECTION – XI

BANK GUARANTEE FORM FOR EMD

Whereas ______________________________ (hereinafter called the “Tenderer”) has submitted its quotation dated ___________________ for the supply of _____________________________ (hereinafter called the “tender”) against the purchaser’s tender enquiry No. __________________________________ Know all persons by these presents that we ______________________________ of _______________________________ (Hereinafter called the “Bank”) having our registered office at ________________________________ are bound unto ____________________________ (hereinafter called the “Purchaser) in the sum of _________________________ for which payment will and truly to be made to the said Purchaser, the Bank binds itself, its successors and assigns by these presents. Sealed with the Common Seal of the said Bank this _____________________ day of _______ 20____. The conditions of this obligation are:

- (1) If the Tenderer withdraws or amends, impairs or derogates from the tender in any respect within the period of validity of this tender.

- (2) If the Tenderer having been notified of the acceptance of his tender by the Purchaser during the period of its validity:-

- a) fails or refuses to furnish the performance security for the due performance of the contract.

- or

- b) fails or refuses to accept/execute the contract.

- or

- c) if it comes to notice that the information/documents furnished in its tender is incorrect, false, misleading or forged

We undertake to pay the Purchaser up to the above amount upon receipt of its first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser will note that the amount claimed by it is due to it owing to the occurrence of one or both the two conditions, specifying the occurred condition(s).

This guarantee will remain in force for a period of forty-five days after the period of tender validity and any demand in respect thereof should reach the Bank not later than the above date.

(Signature of the authorised officer of the Bank)

Name and designation of the officer

Seal, name & address of the Bank and address of the Branch

SECTION – XII

BANK GUARANTEE FORM FOR PERFORMANCE SECURITY/ CAMC SECURITY

To

HLL Lifecare Limited,

Akkulam Factory, Sreekariam PO,

Thiruvananthapuram – 17.

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: hcdcmo@lifecarehll.com

WHEREAS _____________________________ (Name and address of the supplier) (Hereinafter called “the supplier”) has undertaken, in pursuance of contract no________________________ dated _____________ to supply (description of goods and services) (herein after called “the contract”).

AND WHEREAS it has been stipulated by you in the said contract that the supplier shall furnish you with a bank guarantee by a scheduled commercial bank recognised by you for the sum specified therein as security for compliance with its obligations in accordance with the contract;

AND WHEREAS we have agreed to give the supplier such a bank guarantee;

NOW THEREFORE we hereby affirm that we are guarantors and responsible to you, on behalf of the supplier, up to a total of. ________________________ (Amount of the guarantee in words and figures), and we undertake to pay you, upon your first written demand declaring the supplier to be in default under the contract and without cavil or argument, any sum or sums within the limits of (amount of guarantee) as aforesaid, without your needing to prove or to show grounds or reasons for your demand or the sum specified therein.

We hereby waive the necessity of your demanding the said debt from the supplier before presenting us with the demand.

We further agree that no change or addition to or other modification of the terms of the contract to be performed there under or of any of the contract documents which may be made between you and the supplier shall in any way release us from any liability under this guarantee and we hereby waive notice of any such change, addition or modification.

This guarantee shall be valid up to 36 (thirty Six) months from the date of Notification of Award i.e  up to ----------- (indicate date)

…………………………….

(Signature with date of the authorised officer of the Bank)

………………………………………………………….

Name and designation of the officer

………………………………………………………….

………………………………………………………….

Seal, name & address of the Bank and address of the Branch

SECTION – XIII

- MANUFACTURER’S/DISTRIBUTOR’S AUTHORISATION FORM

To

Deputy General Manager (Materials),

HLL Lifecare Limited,

Akkulam Factory, Sreekariam PO,

Thiruvananthapuram – 17.

Phone +91 471 244 5930, Fax  +91 471 244 5935

Email: hcdcmo@lifecarehll.com

Dear Sirs,

Ref. Your TE document No ____________, dated _____________

We, ___________________________________ who are proven and reputable manufacturers/distributor’s of___________________________(name and description of the goods offered in the tender) having factories at_______________________________, hereby authorise Messrs______________________________(name and address of the agent) to submit a tender, process the same further and enter into a contract with you against your requirement as contained in the above referred TE documents for the above goods manufactured by us.

We further confirm that no supplier or firm or individual other than Messrs. __________________________ (name and address of the above agent) is authorised to submit a tender, process the same further and enter into a contract with you against your requirement as contained in the above referred TE documents for the above goods manufactured by us.

We also hereby extend our full warranty, CAMC as applicable as per clause 15 of the General Conditions of Contract, read with modification, if any, in the Special Conditions of Contract for the goods and services offered for supply by the above firm against this TE document.

Yours faithfully,

[Signature with date, name and designation]

for and on behalf of Messrs___________________________

[Name & address of the manufacturer/distributor]

Note :This letter of authorisation should be on the letter head of the manufacturing firm and should be signed by a person competent and having the power of attorney to legally bind the manufacturer.

- CONSORTIUM AGREEMENT

This Consortium agreement (hereinafter referred as “Agreement”) is entered on this the                  __day of ______, 20__ by and between;

M/s ______________, incorporated as a company under the Companies Act, 1956 and having its Registered Office at ___________________ (hereinafter referred as “Lead Partner”), which the expression shall unless otherwise include all its successors, and permitted assigns)  and represented by ____________, in the capacity of ________________ of the FIRST PART

AND

M/s  _____________, a company incorporated under the laws of __________________ and having its Registered Office at _____________________________ (hereinafter referred as “Second Partner”), which the expression shall unless otherwise include all its successors, and permitted assigns)and represented by …….. in the capacity of ________________  of the SECOND PART

For the purpose of this Agreement both the “Lead Partner” as well as “Second Partner” are collectively called “Partners” and individually called “Partner”.

Whereas –

- HLL Lifecare Limited (Purchaser/HLL) has floated a tender vide Tender No: _____________________________ for setting up of Neonatal Intensive Care Unit on turnkey basis.

- As per the Tender document, bids are to be submitted by any Consortium which will be considered; provided such bids fulfill all the specific requirements in that regard.

- Now the Parties to this Agreement decided to form a Consortium to participate in the Tender .

- AND WHEREAS the bid is being submitted based on the consortium agreement being these presents  and the bid with its bid forms and submission documents in accordance with the requirement of tender document conditions and requirements have been signed by all the partners and submitted to HLL Lifecare Ltd.

NOW THIS AGREEMENT WITNESSTH HEREIN AS FOLLOWS

- That the Parties to this Consortium do hereby agrees to participate in the Tender in the name and style of “  _________” (hereinafter referred as “Consortium”).

- Scope: Purpose of this Agreement is to participate and submit all necessary bid documents against the Tender floated by HLL and in case of award, supply the tendered items listed below as against each partner.

| SI. No. | List of Equipments | Qty in Nos |
| --- | --- | --- |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| 6 |  |  |

- Tenure: This Agreement shall be valid till the date of either rejection of the Bid submitted by this Consortium against the Tender floated by the HLL or till the expiry of the Contract entered between the Consortium members and HLL in case of award of the Tender to this Consortium.

- In consideration of the bid submission by us to HLL, pre-qualification of our  technical bid by  HLL  if considered acceptable, submission of price bid by us and the award of contract by HLL to the Consortium (if selected by ), we the partners to the Consortium, hereby agree that                      M/s __________ shall act as the Lead Partner for self, and for and on behalf of Partner – II and further declare and confirm that ______ shall be  solely bound to HLL for execution of the contract in accordance with the contract terms and shall  perform all contractual obligations including technical guarantees. Further, the Lead Partner is authorized to incur liabilities and receive instructions for and on behalf of any or all partners of the Consortium.

- The Lead Partner shall be solely responsible for Management of all the  works to be undertaken under the tender and it shall be the nodal point for HLL for queries, purchase orders, installation and payments.

- In case of any breach of the said Contract by any of the partners of the Consortium, we hereby agree to be fully responsible for the successful execution/ performance of the Contract in accordance with the terms of the Contract.

- Further, if HLL suffers any loss or damage on account of any breach of the Contract or any shortfall in the completed work meeting the guaranteed performance parameters as per the technical specifications/ contract documents, the Lead Partner undertakes to promptly make good such loss or damage caused to HLL, on HLL’s demand without any demure. HLL shall have the right to proceed against_____.

- The financial liability of the partners to this Consortium Agreement, to HLL with respect to the any or all claims arising out of the performance or non-performance of the Contract shall, however be not limited in any way so as to restrict or limit the liabilities of either of the partner.

- It is expressly understood and agreed between the partners to this agreement that the responsibilities and obligations of each of the partners shall be as delineated in this agreement. It is further agreed by the partners that the above sharing of responsibilities and obligations shall not in any way be a limitation of the joint and several responsibilities of the partners under the Contract.

- Obligations of the Second Partner

  - That, the Second Partner ensures the procurement and supply of the items listed above.

  - That, the Second Partner ensures to provide necessary training to the staffs employed in respective training centers operating under the provisions of the contract signed between the Consortium and HLL.

  - That, the Second Partner agrees to provide necessary repairers and replacements for supplied items, if any found defective during the tenure of the agreement between the Consortium and _________HLL Provided such defects have incurred due to any breakage or manufacturing defect and must be pointed out by the Lead Partner to the Second Partner in writing within 3 days from the date of identification of such defect.

- This Consortium Agreement shall be governed, construed and interpreted in accordance with Laws of India. Courts of Thiruvananthapuram shall have exclusive jurisdiction in all matters arising there under.

- In case of award of contract, we the partners to this Consortium Agreement do hereby agree that we shall furnish the contract performance guarantee (if any) in favour of the HLL from a bank acceptable/ approved by  HLL for a value as stipulated in the Contract Award.

- It is further agreed that this Consortium Agreement shall be irrevocable and shall form an integral part of the Contract and shall continue to be enforceable till the Consortium members discharges the same. It shall be effective on the date first above mentioned for all purposes and intents.

IN WITNESS WHEREOF, the partners to this Consortium agreement have, through their respective authorized representatives, have executed these presents and affixed their hands and common seal of their respective companies on the day, month and year first abovementioned.

| For M/s ________________________ Authorized Signatory | For M/s ________________________ Authorized Signatory |
| --- | --- |

Witnesses:

SECTION – XIV

CONTRACT FORM - A

CONTRACT FORM FOR SUPPLY, INSTALLATION, COMMISSIONING, HANDING OVER, TRIAL RUN, TRAINING OF OPERATORS & WARRANTY OF GOODS

THIS AGREEMENT made the .............. day of ................, 20........ between ................... (Name and Address of Purchaser) represented by the General Manager (Mtls)....................... (hereinafter “the Purchaser”) of one part and	(Name and Address of Supplier) ....................................(hereinafter “the Supplier”) represented by……………………… (Name  of  the  Authorized  Signatory and  Designation),  Aged  ….....  years,  residingat ……………………………… (Full Residential Address of the Signatory) of the other part:

WHEREAS the Purchaser is desirous that certain Goods and ancillary services should be provided   by  theSupplier, viz., (Brief Description of Goods andServices)and_______has accepted a bid by the Supplier for the supply of those goods and services in the sum of...................... (Contract Price in Words and Figures) (hereinafter “the Contract Price”).

# NOW THIS AGREEMENT WITNESSETH AS FOLLOWS:

- In this Agreement words and expressions shall have the same meanings as are respectively assigned to them in the Conditions of Contract referred to.

- The following documents shall be deemed to form and be read and constructed as part of this Agreement, viz.:

  - the Bid Form and Price Schedule submitted by the Bidder;

  - the Schedule of Requirements;

  - the Technical Specifications;

  - the General Conditions of Contract;

  - the Special Conditions of Contract ;and

  - the Purchaser’s Notification of Award

- In consideration of the payments to be made by the Purchaser to the Supplier as hereinafter mentioned, the Supplier hereby covenants with the Purchaser to provide the Goods and Services and to remedy defects therein in conformity in all respects with the provisions of the Contract.

- The Purchaser hereby covenants to pay the Supplier in consideration of the provision of the Goods and Services and the remedying of defects therein, the Contract Price or such other sum as may become payable under the provisions of the Contract at the times and in the manner prescribed by the Contract.

# BRIEF PARTICULARS OF THE GOODS AND SERVICES WHICH SHALL BE SUPPORTED / PROVIDED BY THE SUPPLIER ARE:

| SI. No. | Decription of Goods | Qty | Unit Price in Rs. | GST Amt in Rs. | Total Amt in Rs. |
| --- | --- | --- | --- | --- | --- |

Total Value in Rs.______________________________________________

NOW THIS AGREEMENT WITNESSETH AS FOLLOWS:

- In this Agreement words and expressions shall have the same meanings as are respectively assigned to them in the Conditions of Contract referred to.

- The following documents shall be deemed to form and be read and constructed as part of this Agreement, Ref No. viz.:

  - HLL Tender Ref No. ____________________________________

  - HLL Amendment/Corrigendum No.__________________________,

  - Supplier Bid Ref No._____________________________________.

  - Minutes of Negotiation Meeting Dt.__________________________.

  - P.O.Ref  No___________________________________________.

- Delivery location:. THE STATE DRUG TESTING LABORATORY, PANDIT NEHRU OFFICE COMPLEX, GORKHABASTI, KUNJABAN, AGARTHALA, TRIPURA - 799 006

- Completion Period: 60 days from the date of order.

- Terms of Payment: Payment shall be made in Indian Rupees as specified in the contract in the following manner:-

- Payment Terms  -  Equipment Supply

## a)	50% of supply value shall be released against supply and certificate for receipt of the item in good condition from the Hospital/purchaser.

## b)   45% payment shall be released against certificate of installation and commissioning certified by the Hospital Authorities/purchaser.

c)     Final 5% shall be released against submission of certificate issued by Hospital/purchaser certifying that the facility has been installed, commissioned and handed over and submission on Performance Security and other documents stated in this tender.

d)    Payment shall be made subject to recoveries, if any, by way of liquidated damages or any other charges as per terms & conditions of contract and upon the submission of the following documents:

i.  Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount;

ii. Consignee Receipt Certificate as per Section XVII in original issued by the authorized representative of the consignee;

iii. Two copies of packing list identifying contents of each package;

iv.  Inspection certificate issued by the nominated Inspection agency, if any.

v.  Insurance Certificate as per GCC Clause 11 and documents also to be submitted for payment of LC confirming that dispatch documents has already been sent to all concerned as per the contract within 24 hours;

vi.   Certificate of origin.

- Payment Terms  - Turnkey part (Civil & Electrical works)

- 50% of the total value of work may be released upon the completion of 60 % of the work recommended by State In charge (P&S, HLL) and approved by Project Engineer (P&S, HLL)/ Manager (P&S, HLL).

- Balance 50% will be paid along with the final bill only after submission of Work Completion certificate issued by the concerned hospital and approved by Project Engineer (P&S, HLL)/ Manager (P&S, HLL).

- 5% amount of the Bill will be retained as retention money and will be released only after the Defect liability period of 12 months.

- Final Bill to be submitted detailing the work description, quantity and rate as per the Work Order. Payment will be made against actual measurements recorded and certified jointly by State In Charge (P&S, HLL) and the Contractor ’s representative, counter signed by Project Engineer (P&S, HLL)/ Manager (P&S, HLL). The work completion certificate will be issued by the concerned centres/hospitals and it is the responsibility of the contractor to collect the completion certificate certified by the concerned hospital authority & HLL authorities and submit to HLL for balance 50% payment

- The running account bills are to be submitted detailing the work description, quantity and rate as per the Work Order. Payment will be made against actual measurements recorded and certified jointly by State in charge (P&S, HLL) and the Contractor’s representative. For supply of capital items, duly certified delivery challan/supporting documents such as Warranty Certificates etc. shall be enclosed along with bill.

- Tax Deduction: All statutory deductions like GST, Income Tax, Works Contract Tax, E.S.I., P.F. or any other government-imposed liabilities shall be borne by the Contractor (as applicable at the time of execution of job) and shall be deducted from each bill submitted by the Contractor.

- Warranty: Warranty period will be 3 years with effective from the date of installation, commissioning and acceptance. During warranty period 3 preventive maintenance shall be executed, once in every four months in a year.

- Performance Security: The supplier shall submit performance security by way of security deposit or bank guarantee in the form prescribed  for the 5% of the total value of the purchase order, valid upto sixty (60) days after the date of completion of all contractual obligations by the supplier, including the warranty obligations. Necessary preventive maintenance reports shall be submitted at the end of warranty period with working satisfactory certification from HLL/Hospital.

- CAMC: Comprehensive Annual Maintenance Contract as per the prevailing rate as mentioned below will be applicable for the purchase orders issued for Five (5) years after completion of warranty period.

- Purchaser Rights: Purchaser reserves all rights to cancel the purchase order issued by the purchaser at any time without assigning any reasons.

- Applicable Law

The contract shall be governed by and interpreted in accordance with the laws of India for the time being in force.

11. Resolution of disputes

If the parties fail to resolve their dispute or difference by such mutual consultations within twenty one days of its occurrence the same shall be referred  for arbitration  to a Sole arbitrator to be appointed by the C&MD,  of the Purchaser Arbitration shall be conducted as per the provisions of the Arbitration and Conciliation Act 1996 or any statutory modification or re-enactment thereof for the time be in force. The arbitrator shall  give  a reasoned award.

- The award of the arbitrator shall be final and binding on the parties to this contract. The venue of arbitration shall be Thiruvananthapuram, Kerala, India and the language of the proceedings shall be English

- Subject to above mentioned Arbitration Clause, any dispute or differences arising out of this Agreement shall fall under the exclusive jurisdiction of courts at Thiruvananthapuram.

IN WITNESS whereof the parties hereto have caused this Agreement to be executed in accordance with their respective laws the day and year first above written.

Signed, Sealed and Delivered by the

said......................................... (For the Purchaser)

in the presence of ...................................................

Signed, Sealed and Delivered by the

said ...........................................(For the Supplier) (Signature, Name, Designation and

Address with Office seal)

in the presence of ....................................................

- (Signature, Name and Address of witness)

# (Signature, Name and Address of witness

SECTION – XV

CONTRACT FORM – B

CONTRACT FORM FOR COMPREHENSIVE ANNUAL MAINTENANCE CONTRACT

This AGREEMENT made on this day ----------------------- between HLL Lifecare Ltd (HLL)/ THE STATE DRUG TESTING LABORATORY, PANDIT NEHRU OFFICE COMPLEX, GORKHABASTI, KUNJABAN, AGARTHALA, TRIPURA - 799 006, represented by the _______________ having his Office at __________________ (hereinafter called HLL/ THE STATE DRUG TESTING LABORATORY, PANDIT NEHRU OFFICE COMPLEX, GORKHABASTI, KUNJABAN, AGARTHALA, TRIPURA - 799 006) of one part and M/s. _________ represented by --------------- aged --------- years, having his / her Office at --------------------------------------------- (hereinafter called “_______”) of the other part. (The term HLL/ THE STATE DRUG TESTING LABORATORY, PANDIT NEHRU OFFICE COMPLEX, GORKHABASTI, KUNJABAN, AGARTHALA, TRIPURA - 799 006 and ________, wherever the context appear and unless, it is specifically excluded, shall mean and exclude its successor, assign administrators and executors).

WHEREAS the _______ had supplied and installed _____ number ________ at ______ against the supply order placed by HLL vide P.O. No. ______ and as per provisions of the tender and supply order ______ should provide Comprehensive Annual Maintenance Services for this equipment (herein after called services) and the purchaser accepted & approved the rates and terms & conditions offered by __________ in the financial bid for the services to which this agreement made for.

NOW THE AGREEMENT WITNESS AS FOLLOWS:

- In this Agreement, words and expressions shall have the same meanings as are respectively assigned to them in the Conditions of Contract / order referred to above.

- The following documents shall be deemed to form and be read and construed as part of this Agreement, viz:

  - the terms and conditions stipulated in Purchaser’s tender document ref. ________________, Dt._______ for the supply of the equipment.

  - the Bid Form and Price Schedule submitted by ______ for supply and providing the maintenance Services, against the tender.

  - the supply order for supply and installation of the equipment vide Ref No. ____________ Dt.______ placed by Purchaser.

- The Contract of Annual Comprehensive Maintenance is hereby concluded as under:

| Sr. No | Name of Equipment | Qty | Total CAMC Amt in Rs. (Excl applicable GST) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 6th year | 7th year | 8th year | 9th year | 10th year |

- Payments shall be made by HLL/ THE STATE DRUG TESTING LABORATORY, PANDIT NEHRU OFFICE COMPLEX, GORKHABASTI, KUNJABAN, AGARTHALA, TRIPURA - 799 006 to Supplier as per the Purchaser’s tender document, _______ hereby covenant with the Purchaser to provide the Comprehensive Annual Maintenance Services in conformity in all respects with the provisions of the Contract and the orders referred above.

- The rates indicated cover all charges towards cost of spare parts, transportation and installation charges, cost of travel, boarding, lodging and expenses related to service personnel and other expenses related to maintenance of the equipment. No claim whatsoever will be entertained.

- It is agreed that the rates indicated hereunder will be firm during the contract period and the contract period is for 5 years from ________.(date of expiry of 5 year warranty period)

- Performance security shall be submitted by way of Bank Guarantee valid till	 [(fill the date) 2 months after expiry of entire CAMC period] for an amount of Rs._________ [(fill amount) equivalent to 5 % of the cost of the total CAMC value.

- It is agreed that _______ will provide preventive maintenance call atleast one visit in four months and the gap between any two Preventive Maintenance should not be less than 90 days and not more than 115 days and attend all breakdown calls, within the time limit prescribed in the tender. In addition, all breakdown calls shall be attended  to  immediately  and  all  major repairs shall be rectified within 7 calendar days from the date of intimation,  as per tender, failing which, a penalty of Rs.50,000/- per day thereof is leviable until the equipment is repaired and commissioned to the    satisfaction  of   the  end  user,  For  this  purpose  you  shall  carry   sufficient   inventories to assure prompt replacement of defective parts as per tender.

- It is agreed that the failure to attend to any breakdown calls within the prescribed time limit will attract penalty as stipulated in the CAMC order.

- Uptime guarantee of 95% shall be maintained by _______ on annual basis taking into consideration the number of actual working hours and working days of the centre.

- Purchaser reserves its rights to get the maintenance services done through any other agency at your full risk and cost and also to take appropriate penal action including termination of the contract, if the performance of services is found not satisfactory and below the 95% uptime level.

- All disputes arising out of this agreement would be settled by  arbitration by a  sole arbitrator to be appointed by the CMD Of  Purchaser

- All disputes arising out of this agreement will be subjected to the jurisdiction of Thiruvananthapuram only.

IN WITNESS whereof the parties hereto have caused this Agreement to be executed in accordance with their respective laws the day and year first above written.

Signed, Sealed and Delivered by the

Said ……………………………………………….. (For the THE STATE DRUG TESTING LABORATORY, PANDIT NEHRU OFFICE COMPLEX, GORKHABASTI, KUNJABAN, AGARTHALA, TRIPURA - 799 006)

in the presence of ……………………………………………..

Signed, Sealed and Delivered by the

Said ……………………………………………… (For __________)

in the presence of …………………………………

SECTION – XVI

CONSIGNEE RECEIPT CERTIFICATE

(To be given by consignee’s authorized representative)

The following store (s) has/have been received in good condition:

- Contract No. & date	:______________________________

- Supplier’s Name	:______________________________

- Consignee’s Name & Address with telephone No. & Fax No. 	: ______________________________

- Name of the item supplied	:______________________________

- Quantity Supplied	:______________________________

- Date of Receipt by the Consignee:______________________________

- Name and designation of Authorized Representative of Consignee	:______________________________

- Signature of Authorized Representative of Consignee with date	:______________________________

- Seal of the Consignee	:______________________________

SECTION – XVII

Proforma of Final Acceptance Certificate by the Consignee

## No_______________

## Date_______________

#### To

M/s	_______________________

_________________________

____________________________

Subject:	Certificate of commissioning of equipment/plant.

This is to certify that the equipment(s)/plant(s) as detailed below has/have been received in good conditions along with all the standard and special accessories and a set of spares (subject to remarks in Para no.02) in accordance with the contract/technical specifications.  The same has been installed and commissioned.

(a)	Contract No______________________________________ dated__________________

(b)	Description of the equipment(s)/plants: _____________________________________

(c)	Equipment(s)/ plant(s) nos.:_______________________________________________

(d)	Quantity:______________________________________________________________

(e)	Bill of Loading/Air Way Bill/Railway

Receipt/ Goods Consignment Note no_______________ dated __________________

(f)	Name of the vessel/Transporter:___________________________________________

(g)	Name of the Consignee:____________________________________________________________

(h)	Date of commissioning and proving test:_____________________________________

Details of accessories/spares not yet supplied and recoveries to be made on that account.

Sl.	Description of Item	Quantity	Amount to be recovered No.

No.

The proving test has been done to our entire satisfaction and operators have been trained to operate the equipment(s)/plant(s).

The supplier has fulfilled its contractual obligations satisfactorily ## or

The supplier has failed to fulfil its contractual obligations with regard to the following:

He has not adhered to the time schedule specified in the contract in dispatching the documents/drawings pursuant to ‘Technical Specifications’.

He has not supervised the commissioning of the equipment(s)/plant(s)in time, i.e. within the period specified in the contract from date of intimation by the Purchaser in respect of the installation of the equipment(s)/plant(s).

The supplier as specified in the contract has not done training of personnel.

The extent of delay for each of the activities to be performed by the supplier in terms of the contract is

The amount of recovery on account of non-supply of accessories and spares is given under Para no.02.

The amount of recovery on account of failure of the supplier to meet his contractual obligations is__________________ (here indicate the amount).

Signature.

Name.

Designation with stamp.

## Explanatory notes for filling up the certificate:

He has adhered to the time schedule specified in the contract in dispatching the documents/drawings pursuant to ‘Technical Specification’.

He has supervised the commissioning of the equipment(s)/plant(s) in time, i.e. within the time specified in the contract from date of intimation by the Purchaser in respect of the installation of the equipment(s)/plant(s).

#### Training of personnel has been done by the supplier as specified in the contract.

In the event of documents/drawings having not been supplied or installation and commissioning of the equipment(s)/plant(s) having been delayed on account of the supplier, the extent of delay should always be mentioned.

- SECTION XVIII

- CONSIGNEE ADDRESS

- STATE DRUG TESTING LABORATORY

- PANDIT NEHRU OFFICE COMPLEX

- GORKHABASTI, KUNJABAN

- AGARTHALA,

- TRIPURA 799 006.

SECTION – XIX

CHECK LIST

### Name of Tenderer:

### Name of Manufacturer:

| Sl No. | Activity | Yes/ No/ NA | Page No. in the TE document | Remarks |
| --- | --- | --- | --- | --- |
| 1. a. | Have you enclosed EMD of required amount for the quoted schedules? |  |  |  |
| b. | In case EMD is furnished in the form of Bank Guarantee, has it been furnished as per Section XIII? |  |  |  |
| c. | In case Bank Guarantee is furnished, have you kept its validity of 165 days from Techno Commercial Tender Opening date as per clause 19 of GIT? |  |  |  |
| 2. a. | Have you enclosed duly filled Tender Form as per format in Section IX? |  |  |  |
| b. | Have you enclosed Power of Attorney in favour of the signatory? |  |  |  |
| 3. a. | Have you enclosed clause-by-clause technical compliance statement for the quoted goods vis-à-vis the Technical specifications? |  |  |  |
| b. | In case of Technical deviations in the compliance statement, have you identified and marked the deviations? |  |  |  |
| 4. a. | Have you submitted satisfactory performance certificate as per the Proforma for performance statement in Sec. IX of TE document in respect of all orders? |  |  |  |
| b. | Have you submitted copy of the order(s) and end user certificate? |  |  |  |
| 5. | Have you submitted manufacturer’s / distributors authorization as per Section XIV? |  |  |  |
| 6. | Have you submitted prices of goods, CAMC etc. in the Price Schedule as per Section X? |  |  |  |
| 7. | Have you kept validity of 120 days from the Techno Commercial Tender Opening date as per the TE document? |  |  |  |
| 8. a. | In case of Indian Tenderer, have you furnished Income Tax Account No. as allotted by the Income Tax Department of Government of India? |  |  |  |
| b. | In case of Foreign Tenderer, have you furnished Income Tax Account No. of your Indian Agent as allotted by the Income Tax Department of Government of India? |  |  |  |
| 9. | Have you intimated the name and full address of your Banker (s) along with your Account Number |  |  |  |
| 10. | Have you fully accepted payment terms as per TE document? |  |  |  |
| 11. | Have you fully accepted delivery period as per TE document? |  |  |  |
| 12. | Have you submitted the certificate of incorporation? |  |  |  |
| 13. | Have you accepted the warranty as per TE document? |  |  |  |
| 14. | Have you accepted terms and conditions of TE document? |  |  |  |
| 15. | Have you furnished documents establishing your eligibility & qualification criteria as per TE documents? |  |  |  |
| 16. | Have you furnished relevant pages of Annual Report (Balance Sheet and Profit & Loss Account) regarding turnover and profit for last three years prior to the date of Tender opening? |  |  |  |

Date:

Name

Signature

Stamp and full address