Date: 18/06/2015

Amendment No: 1

Ref: Tender Enquiry No: HLL/BME/Calibration-06/15-16       Dated 11.06.2015

Pre-Bid meeting against the referred tender was held on 18/06/2015 at  HLL Lifecare Ltd., Thycaud, Trivandrum- 695 014

The following amendments are issued with respect to the above Tender Enquiry Document.

I )Technical Specification

Item No 3

Power and Energy Analyzer

Point 2.2

Existing Specification: Current Measurement Range (ac+dc): 5A to 600A

Amended as: Current Measurement Range (ac+dc): 0.5A to 600A

Following lines are added to the existing specification:

Point 2.4: Provision for Invertor efficiency is desirable.

Item No 4

Endoscope Tester

Point 3.1

Existing Specification:Comprehensive warranty for 2 years

Amended as: deleted

Item No 6

Data logger temperature

Point 3.1

Existing Specification:Data logger with at least 12 channels- 01

Amended as: Data logger with at least 10 channels and upgradable- 01

Point 4.2

Existing Specification:During the warranty period, the equipment should be calibrated free of cost

Amended as:deleted

Item No 10

Spectrum Analyzer

Point1

Existing Specification: Should be a spectrum analyser to measure occupied Bandwidth, Channel Power, ACPR, C/I, Field Strength, Spectral Emissions, Interference like Spectrogram, Signal Strength, RSS etc.

Amended as: Should be a spectrum analyser to measure occupied Bandwidth, Channel Power, ACPR, C/I, Field Strength, Spurious Emissions, Signal Strength etc.

Point 2.1

Existing Specification: Frequency: 9 KHz to 3GHz

Amended as:Frequency: 9 KHz to 9GHz

Point 2.10

Existing Specification: should to be able to plot RSSI on on-screen map

Amended as:deleted

Point 2.11

Existing Specification: Should have Interference Mapping

Amended as:deleted

Item No 15

Laser Power Meter

Point 2.5

Existing Specification:The device should have peltier based thermal head  with wavelength range 180  to 20,000 and input power range 1 mW to 20 W

Amended as:Deleted

Point 3.4

Existing Specification: Thermal detector head - 01

Amended as: Deleted

(II)Section VI  List of Requirements

Amended as

| Sl. No. | Item | Warranty period (yrs.) | CMC period (yrs.) |
| --- | --- | --- | --- |
| 1 | Dialysis Reference Meter | 2 | 5 |
| 2 | Phototherapy Radiometer | 2 | 5 |
| 3 | Power Analyzer | 2 | 5 |
| 4 | Endoscope Tester | Nil | Nil |
| 5 | CR/DR Test tool | Nil | Nil |
| 6 | Data Logger | 2 | 5 |
| 7 | Frequency Counter | 2 | 5 |
| 8 | Earth Tester | 2 | 5 |
| 9 | Ultrasound Leakage Tester | 2 | 5 |
| 10 | Spectrum analyzer | 2 | 5 |
| 11 | Ultrasound Phantom | Nil | Nil |
| 12 | Mammography Phantom | Nil | Nil |
| 13 | CT Imaging Phantom | Nil | Nil |
| 14 | High & Low Contrast Tool(s) | Nil | Nil |
| 15 | Laser Power meter | 2 | 5 |

(III)Service Manuals should be provided with Item no 1 Dialysis Reference meter, Item no 6 Data Logger, Item no 7 Frequency Counter, Item no 8 Earth Tester, Item no 9 Ultrasound Leakage Tester, Item no 15 Laser Power meter.

All other contents, terms & conditions of the tender enquiry remain unchanged.