Amendment / Corrigendum to Tender

Sub:    TENDER FOR ESTABLISHMENT OF DEIC

IN THE STATE OF HARYANA ON TURNKEY BASIS.

Ref:  IFB NO: HLL/AFT-CMO/HR-DEIC/EQP/2017-18 DATED 14.02.2018

| Sl No |  | Instead of |  | # Read as |  |
| --- | --- | --- | --- | --- | --- |
| 1. |  | RET cam or any other Camera to take photographs of the fundus of the new-born |  | Portable Modular Retinal Camera for Premature Babies (ROP) |  |
| CORRIGENDUM / AMENDMENT TO SECTION VI-Technical Specification (Page – 111) |  |  |  |  |  |
| SPECIFICATIONS PORTABLE MODULAR RETINAL CAMERA FOR PREMATURE BABIES( ROP Device) - It should be Light weight and portable modular retinal camera for diagnosis of retinal diseases of the premature babies. - Minimum 120 degree field of view. - Should be able to capture images of the entire eye in a video mode and should be convertible to an image mode for email. - Retinal camera should be tele medicine compatible so that the high quality and high resolution images of the infant eye could be shared with anyone anywhere for views. - Software should support sending and receiving images and diagnosis. - Should have high quality image editing tool to convert the film clip of the screening to frames and freedom of selecting frames. - Should also assist in image enhancement and compression for faster and efficient transmission. - Should have adjustable intensity, gain, balance, brightness, contrast, and gamma. - Should be compatible with DICOM export, Telemedicine, USB 3.0, MS WINDOW 8.1 or better. - Image resolution ------ 2040x2040-24 bit colors. - Should be able to measure pupil size of 4mm and above. - Minimum 120 degree field of view from center of the eye. - Should supplied with branded laptop with following specification - High Speed i5 Processor or higher, RAM 8GB or more, Hard disk -1 TB or more, Window 8.1 or better, Display at least 15 inch or more |  |  |  |  |  |
| Sl No |  |  | Original Tender Details |  | # Amended Tender Details |  |
| 1. |  | Bid Submission Closing Date 28/02/2018 at 14.00 hrs Bid Opening Date 28/02/2018 at 15.00 hrs |  | Bid Submission Closing Date 05/04/2018 at 14.00 hrs Bid Opening Date 05/04/2018 at 15.00 hrs |  |

All other terms and conditions in the original tender remain unchanged.

21-03.2018                                                                        Joint General Manager (Materials)